## PID Summary
Part1: PID Summary | Part1: PID Summary | Part1: PID Summary | Part1: PID Summary | Part1: PID Summary
JobNo | GPN | QTY | Description | Job_Date
16045256 | 012-04524-FH | 540 | PCB Assy, Odin / Modi, WHR Flex+WHR Board | 2022/09/21
| Total | 540
Test Result
Process Type | ItemNameType | FPY % | FYR % | Retest Rate % | Retest Rate(WO. ProductFail) % | Tested ESN QTY / Output | Avg Pass (sec) | TestDate (UTC)
Epix_Pro_ECG_ATE | 17858 | 91.11 | 97.96 | 12.41 | 12.41 | 540/529 | 126.122 | 2022-09-22T10:23:51.407+00:00
Epix_Pro_ATE_TI | 17859 | 68.24 | 98.49 | 55.77 | 55.77 | 529/521 | 146.2371 | 2022-09-23T03:40:59.313+00:00

## Test Result
Part 3: Failure Distribution
PPG ATE
Failitemname | Count | Fail Rate | Analyzed | Uncertain
Barometer Temperature | 5 | 0.9452%
AFE Power Consumption data | 3 | 0.5671%
D104 PD dark noise OHR dev | 2 | 0.3781%
D101 PD dark noise SpO2 dev | 1 | 0.189%
D102 PD dark noise OHR dev | 1 | 0.189%
D103 PD dark noise SpO2 dev | 1 | 0.189%
D103 Response avg | 1 | 0.189%
D110 ?V 940nm Brightness (5mA) Measure diff with V_ref | 1 | 0.189%
Temperature Sensor Temperature | 1 | 0.189%
AFE Part ID Read | 1 | 0.189%
Failure list
ESN | GPN | FailItemName | FailLocation | FailReason | JOBNO | Comment | Photo
2905E04139 | 012-04524-FH | D103 Response avg
2905E04594 | 012-04524-FH | Barometer Temperature
2905E04661 | 012-04524-FH | Barometer Temperature
2905E04960 | 012-04524-FH | Barometer Temperature
2905E03579 | 012-04524-FH | Barometer Temperature
2905E03705 | 012-04524-FH | Barometer Temperature
2905E04866 | 012-04524-FH | Temperature Sensor Temperature
2905E04661 | 012-04524-FH | AFE Part ID Read
2905E04661 | 012-04524-FH | D101 PD dark noise SpO2 dev
2905E04661 | 012-04524-FH | D103 PD dark noise SpO2 dev
2905E04594 | 012-04524-FH | D102 PD dark noise OHR dev
2905E04594 | 012-04524-FH | D104 PD dark noise OHR dev
2905E04661 | 012-04524-FH | D104 PD dark noise OHR dev
2905E04594 | 012-04524-FH | D110 ?V 940nm Brightness (5mA) Measure diff with V_ref
2905E03705 | 012-04524-FH | AFE Power Consumption data
2905E04842 | 012-04524-FH | AFE Power Consumption data
2905E04960 | 012-04524-FH | AFE Power Consumption data
ECG ATE
Failitemname | Count | Fail Rate | Analyzed | Uncertain
Motor Resonant Frequency | 7 | 1.296%
Check 1907IOP | 2 | 0.3704%
ECG 2mv 5Hz Test | 2 | 0.3704%
ECG 2mv 1Hz Test | 1 | 0.1852%
ECG 2mv 2Hz Test | 1 | 0.1852%
Is ESN check in | 1 | 0.1852%
Failure list
ESN | GPN | FailItemName | FailLocation | FailReason | JOBNO | Comment | Photo
2905E04159 | 012-04524-FH | Is ESN check in
2905E04904 | 012-04524-FH | Check 1907IOP
2905E04600 | 012-04524-FH | Check 1907IOP
2905E04597 | 012-04524-FH | ECG 2mv 5Hz Test
2905E04159 | 012-04524-FH | ECG 2mv 5Hz Test
2905E04159 | 012-04524-FH | ECG 2mv 1Hz Test
2905E03774 | 012-04524-FH | Motor Resonant Frequency
2905E04045 | 012-04524-FH | Motor Resonant Frequency
2905E05225 | 012-04524-FH | Motor Resonant Frequency
2905E05007 | 012-04524-FH | Motor Resonant Frequency
2905E05053 | 012-04524-FH | Motor Resonant Frequency
2905E04446 | 012-04524-FH | Motor Resonant Frequency
2905E04030 | 012-04524-FH | Motor Resonant Frequency
2905E04597 | 012-04524-FH | ECG 2mv 2Hz Test
Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate
PPG ATE | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate
ItemName Type | Process Type | Item | ItemName | Count/Count ESN | Retest% | Comment
17859 | Epix_Pro_ATE_TI | 88 | AFE Power Consumption data | 128/529 | 24.2
17859 | Epix_Pro_ATE_TI | 41 | Barometer Temperature | 93/529 | 17.58
17859 | Epix_Pro_ATE_TI | 46 | Temperature Sensor Temperature | 17/529 | 3.21
17859 | Epix_Pro_ATE_TI | 10 | D103 Response avg | 7/529 | 1.32
17859 | Epix_Pro_ATE_TI | 49 | AFE Part ID Read | 5/529 | 0.95
17859 | Epix_Pro_ATE_TI | 71 | D104 PD dark noise OHR dev | 5/529 | 0.95
17859 | Epix_Pro_ATE_TI | 68 | D101 PD dark noise OHR dev | 4/529 | 0.76
17859 | Epix_Pro_ATE_TI | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | 4/529 | 0.76
17859 | Epix_Pro_ATE_TI | 3 | Fixture ID | 3/529 | 0.57
17859 | Epix_Pro_ATE_TI | 4 | USB Connection 1 | 3/529 | 0.57
17859 | Epix_Pro_ATE_TI | 34 | AFE Power Consumption Test | 3/529 | 0.57
17859 | Epix_Pro_ATE_TI | 59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | 3/529 | 0.57
17859 | Epix_Pro_ATE_TI | 69 | D102 PD dark noise OHR dev | 3/529 | 0.57
17859 | Epix_Pro_ATE_TI | 55 | D106 – 540nm Brightness (20mA) Measure diff with V_ref | 2/529 | 0.38
17859 | Epix_Pro_ATE_TI | 70 | D103 PD dark noise OHR dev | 2/529 | 0.38
17859 | Epix_Pro_ATE_TI | 73 | D109 – 630nm Brightness (20mA) Measure diff with V_ref | 2/529 | 0.38
17859 | Epix_Pro_ATE_TI | 76 | D110 – 540nm Brightness (20mA) Measure diff with V_ref | 2/529 | 0.38
17859 | Epix_Pro_ATE_TI | 77 | D110 – 630nm Brightness (20mA) Measure diff with V_ref | 2/529 | 0.38
17859 | Epix_Pro_ATE_TI | 79 | D110 – 940nm Brightness (20mA) Measure diff with V_ref | 2/529 | 0.38
17859 | Epix_Pro_ATE_TI | 21 | D109 (540nm) Brightness and LED Current (20mA) | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 25 | D109 (660nm) Brightness and LED Current (20mA) | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 60 | D101 PD dark noise SpO2 dev | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 62 | D103 PD dark noise SpO2 dev | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 74 | D109 – 660nm Brightness (20mA) Measure diff with V_ref | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 75 | D109 – 940nm Brightness (20mA) Measure diff with V_ref | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 78 | D110 – 660nm Brightness (20mA) Measure diff with V_ref | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 80 | D110 – 940nm Brightness (5mA) Measure diff with V_ref | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 98 | D109 – 540nm Brightness (20mA) | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 104 | D110 – 630nm Brightness (20mA) | 1/529 | 0.19
17859 | Epix_Pro_ATE_TI | 112 | (LED rotation check) Internal Light Pollution (ILP) Ch0 | 1/529 | 0.19
ECG ATE | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate | Part 4: Testing Retry Rate
ItemName Type | Process Type | Item | ItemName | Count/Count ESN | Retest% | Comment
17858 | Epix_Pro_ECG_ATE | 30 | Motor Resonant Frequency | 55/540 | 10.19
17858 | Epix_Pro_ECG_ATE | 20 | ECG 2mv 5Hz Test | 7/540 | 1.3
17858 | Epix_Pro_ECG_ATE | 15 | Check 1907IOP | 5/540 | 0.93
17858 | Epix_Pro_ECG_ATE | 9 | Is ESN check in | 3/540 | 0.56
17858 | Epix_Pro_ECG_ATE | 59 | ECG 2mv 2Hz Test | 2/540 | 0.37
17858 | Epix_Pro_ECG_ATE | 7 | USB Connection (1) | 1/540 | 0.19
17858 | Epix_Pro_ECG_ATE | 27 | ECG 2mv 1Hz Test | 1/540 | 0.19

## PR Issue-Ongoing (hidden)
Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue
Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing
Priority | Status | GPN | Part GPN | Phase | Issue Type | Description | Start Date | Due Date | Department | Owners | Comment | MP Block

## PR Issue-Pending (hidden)
Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue
Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending
Priority | Status | GPN | Part GPN | Phase | Issue Type | Description | Start Date | Due Date | Department | Owners | Comment | MP Block

## PR Issue-Closed (hidden)
Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue | Part2: PR Issue
Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed
Priority | Status | GPN | Part GPN | Phase | Issue Type | Description | Start Date | Due Date | Department | Owners | Comment | MP Block

## Failure Analysis (hidden)
Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis | Part3 : Failure Analysis
Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed | Closed
Phase Name | Status | GPN | Job No | Serial Number | Process Type | Failitem | FailItemName | Fail Value | Start Date | Department | Owner | Ship to | Uploaded Files
Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing | Ongoing
Phase Name | Status | GPN | Job No | Serial Number | Process Type | Failitem | FailItemName | Fail Value | Start Date | Department | Owner | Ship to | Uploaded Files
Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending | Pending
Phase Name | Status | GPN | Job No | Serial Number | Process Type | Failitem | FailItemName | Fail Value | Start Date | Department | Owner | Ship to | Uploaded Files

## PPG Test Item
Code | Name | Count | Unit | USL | LSL | MaxValue | MinValue | AvgValue | AvgSpare | MaxSpare | MinSpare | AvgRetries | MaxRetries | MinRetries | Stddev | Ck | Cp | Cpk
1 | Input ESN | 824 (824/0) |  |  |  | 0 | 0 | 0 | 9593.807 | 122198 | 2279 | 0 | 0 | 0 | 0 | - | - | -
2 | Final Index 32767 | 0 (0/0) |  | 32767 | 32767 | - | - | - | - | - | - | - | - | - | - | - | - | -
3 | Fixture ID | 821 (821/0) |  |  |  | 4 | 1 | 2.4409 | 1159.838 | 1313 | 1104 | 0 | 0 | 0 | 1.1909 | - | - | -
4 | USB Connection 1 | 819 (819/0) |  |  |  | 1 | 1 | 1 | 15310.6838 | 180492 | 3389 | 1.475 | 32 | 0 | 0 | - | - | -
5 |  | 540 (540/0) |  |  |  | 21205.27 | 18283.63 | 19928.5393 | 4298.4778 | 4703 | 4117 | 0 | 0 | 0 | 440.4603 | - | - | -
6 |  | 547 (547/0) |  |  |  | 21716.02 | 18773.46 | 20053.8645 | 4298.7861 | 4770 | 4139 | 0 | 0 | 0 | 446.4332 | - | - | -
7 |  | 547 (547/0) |  |  |  | 1523.708 | 988.3154 | 1242.9938 | 1164.7587 | 1428 | 1099 | 0 | 0 | 0 | 89.848 | - | - | -
8 |  | 547 (547/0) |  |  |  | 21133.44 | 18381.74 | 19956.1173 | 4279.4808 | 6502 | 4137 | 0.0018282 | 1 | 0 | 451.9708 | - | - | -
9 |  | 547 (547/0) |  |  |  | 1484.111 | 1016.622 | 1226.1317 | 1164.2011 | 1369 | 1106 | 0 | 0 | 0 | 88.793 | - | - | -
10 |  | 540 (540/0) |  |  |  | 21235.45 | 18489.93 | 19871.233 | 4306.0778 | 6513 | 4147 | 0.0018519 | 1 | 0 | 457.0137 | - | - | -
11 |  | 540 (540/0) |  |  |  | 1444.097 | 997.282 | 1203.1664 | 1166.0056 | 1509 | 1096 | 0 | 0 | 0 | 83.0763 | - | - | -
12 |  | 540 (540/0) |  |  |  | 1704.659 | 933.2512 | 1210.6932 | 1165.5222 | 1430 | 1089 | 0 | 0 | 0 | 88.097 | - | - | -
13 | D105 – 540nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
14 | D106 – 540nm Brightness (20mA) Measure | 0 (0/0) | V | 3.81 | 3.8 | - | - | - | - | - | - | - | - | - | - | - | - | -
15 | D106 (540nm) Brightness and LED Current (20mA) | 570 (570/0) | mA | 22 | 18 | 21.232 | 20.336 | 20.8577 | 1139.6281 | 1590 | 1091 | 0 | 0 | 0 | 0.20044 | -0.42886 | 3.326 | 1.8996
16 | D107 – 540nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
17 | D107 (540nm) Brightness and LED Current (20mA) | 570 (570/0) | mA | 22 | 18 | 21.232 | 20.336 | 20.8587 | 1138.6228 | 1334 | 1088 | 0 | 0 | 0 | 0.20047 | -0.42937 | 3.3256 | 1.8977
18 | D108 – 540nm Brightness (20mA) Measure | 0 (0/0) | V | 3.8 | 3.75 | - | - | - | - | - | - | - | - | - | - | - | - | -
19 | D108 (540nm) Brightness and LED Current (20mA) | 567 (567/0) | mA | 22 | 18 | 21.232 | 20.338 | 20.8611 | 1138.9947 | 1460 | 1068 | 0 | 0 | 0 | 0.19967 | -0.43053 | 3.3389 | 1.9014
20 | D109 – 540nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
21 | D109 (540nm) Brightness and LED Current (20mA) | 561 (561/0) | mA | 22 | 18 | 21.233 | 20.338 | 20.8618 | 1138.1105 | 1416 | 1075 | 0 | 0 | 0 | 0.19952 | -0.4309 | 3.3413 | 1.9015
22 | D109 – 630nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
23 | D109 (630nm) Brightness and LED Current (20mA) | 559 (559/0) | mA | 22 | 18 | 21.234 | 20.337 | 20.8619 | 1137.7567 | 1240 | 1081 | 0 | 0 | 0 | 0.19971 | -0.43095 | 3.3382 | 1.8996
24 | D109 – 660nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
25 | D109 (660nm) Brightness and LED Current (20mA) | 557 (557/0) | mA | 22 | 18 | 21.234 | 20.337 | 20.8627 | 1138.4937 | 1401 | 1088 | 0 | 0 | 0 | 0.19996 | -0.43137 | 3.334 | 1.8958
26 | D109 – 940nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
27 | D109 (940nm) Brightness and LED Current (20mA) | 556 (556/0) | mA | 22 | 18 | 21.238 | 20.343 | 20.8675 | 1138.8129 | 1299 | 1080 | 0 | 0 | 0 | 0.20007 | -0.43374 | 3.3322 | 1.8869
28 | D110 – 540nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
29 | D110 (540nm) Brightness and LED Current (20mA) | 554 (554/0) | mA | 22 | 18 | 21.235 | 20.339 | 20.863 | 1137.7329 | 1319 | 1088 | 0 | 0 | 0 | 0.20018 | -0.43151 | 3.3303 | 1.8932
30 | D110 – 630nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
31 | D110 (630nm) Brightness and LED Current (20mA) | 551 (551/0) | mA | 22 | 18 | 21.236 | 20.34 | 20.8639 | 1139.3521 | 1619 | 1090 | 0 | 0 | 0 | 0.19959 | -0.43193 | 3.3401 | 1.8974
32 | D110 – 660nm Brightness (20mA) Measure | 0 (0/0) | V | 4 | 3.5 | - | - | - | - | - | - | - | - | - | - | - | - | -
33 | D110 (660nm) Brightness and LED Current (20mA) | 550 (550/0) | mA | 22 | 18 | 21.234 | 20.34 | 20.8637 | 1139.4582 | 1391 | 1089 | 0 | 0 | 0 | 0.19956 | -0.43186 | 3.3407 | 1.898
34 | AFE Power Consumption Test | 572 (572/0) | mA | 0.16 | 0.14 | 0.16 | 0.14 | 0.14857 | 1170.8654 | 3448 | 1086 | 0.012238 | 1 | 0 | 0.0038848 | 0.14336 | 0.85804 | 0.73504
35 |  | 0 (0/0) |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | -
36 | AFE Power Consumption Test sampling rate to 16Hz | 575 (575/0) |  | 16 | 16 | 16 | 16 | 16 | 144.6835 | 279 | 109 | 0 | 0 | 0 | 0 | - | - | -
37 |  | 548 (548/0) |  |  |  | 21.242 | 20.346 | 20.8685 | 1139.0091 | 1309 | 1085 | 0 | 0 | 0 | 0.19989 | - | - | -
38 |  | 0 (0/0) |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | -
39 |  | 547 (547/0) |  |  |  | 5.246 | 5.019 | 5.1529 | 1139.1353 | 1523 | 1088 | 0 | 0 | 0 | 0.049075 | - | - | -
40 | AFE Part ID Read | 0 (0/0) |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | -
41 | Barometer Temperature | 725 (725/0) |  | 32 | 18 | 25.2109 | 21.2109 | 24.1753 | 617.6648 | 14116 | 115 | 0.068966 | 12 | 0 | 0.68898 | 0.11781 | 3.3867 | 2.9877
42 | Barometer Pressure | 725 (725/0) |  | 110000 | 90000 | 101038.9 | 100581.6 | 100800.8579 | 12.2841 | 108 | 0 | 0 | 0 | 0 | 117.8075 | -0.080086 | 28.2948 | 26.0287
43 | Query Golden Baro Value | 725 (725/0) |  | 110000 | 90000 | 101037 | 100545 | 100744.0428 | 33.4386 | 601 | 9 | 0 | 0 | 0 | 149.1643 | -0.074404 | 22.3467 | 20.684
44 | CompareBaroPressure | 725 (725/0) |  | 500 | -500 | 157.3125 | -130.125 | 56.8106 | 15.771 | 79 | 0 | 0 | 0 | 0 | 59.6007 | -0.11362 | 2.7964 | 2.4787
45 | Get Baro Pressure | 725 (725/0) |  |  |  | 101038.9 | 100581.6 | 100800.8579 | 45.4979 | 930 | 0 | 0 | 0 | 0 | 117.8075 | - | - | -
46 | Temperature Sensor Temperature | 708 (708/0) |  |  |  | 26.035 | 21.76 | 24.9106 | 373.2585 | 1211 | 45 | 0.035311 | 1 | 0 | 0.75102 | - | - | -
47 | Temperature Sensor status | 708 (708/0) |  | 0 | 0 | 0 | 0 | 0 | 20.8559 | 74 | 9 | 0 | 0 | 0 | 0 | - | - | -
48 | VLED Power Consumption Test | 572 (572/0) | mA |  |  | 0.274 | 0.161 | 0.20093 | 1136.6958 | 1239 | 1077 | 0 | 0 | 0 | 0.022314 | - | - | -
49 | AFE Part ID Read | 703 (703/0) |  |  |  | 808800 | -999 | 769667.441 | 2634.9929 | 50435 | 49 | 0.90612 | 20 | 0 | 173584.8928 | - | - | -
50 | V_ref (measure) | 0 (0/0) | V |  |  | - | - | - | - | - | - | - | - | - | - | - | - | -
51 |  | 540 (540/0) |  |  |  | 26.5981 | -11.7711 | 6.287 | 20.8685 | 89 | 9 | 0 | 0 | 0 | 6.1738 | - | - | -
52 | V_ref | 708 (708/0) | V |  |  | 4.5196 | 4.4573 | 4.4821 | 1236.065 | 2302 | 1154 | 0 | 0 | 0 | 0.023767 | - | - | -
53 | D105 – 540nm Brightness (20mA) Measure diff with V_ref | 572 (572/0) | mV |  | 747 | 946.502 | 858.556 | 898.1399 | 3771.264 | 12025 | 1220 | 1.0857 | 4 | 0 | 16.2716 | - | 3.0962 | 3.0962
54 |  | 540 (540/0) |  |  |  | 29.0756 | -10.9159 | 6.1092 | 2043.6407 | 2228 | 1936 | 1 | 1 | 1 | 6.4295 | - | - | -
55 | D106 – 540nm Brightness (20mA) Measure diff with V_ref | 570 (570/0) | mV |  | 747 | 975.864 | 855.174 | 923.7751 | 1708.793 | 9036 | 1147 | 0.19474 | 3 | 0 | 15.6633 | - | 3.762 | 3.762
56 |  | 540 (540/0) |  |  |  | 28.1827 | -16.1475 | 5.2147 | 20.6315 | 100 | 9 | 0 | 0 | 0 | 6.9512 | - | - | -
57 |  | 540 (540/0) |  |  |  | 36.0552 | -17.1913 | 7.2344 | 20.1963 | 108 | 9 | 0 | 0 | 0 | 10.4529 | - | - | -
58 | D107 – 540nm Brightness (20mA) Measure diff with V_ref | 570 (570/0) | mV |  | 747 | 951.368 | 842.355 | 901.1178 | 3619.2316 | 11868 | 1222 | 1.0281 | 4 | 0 | 15.2155 | - | 3.3763 | 3.3763
59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | 567 (567/0) | mV |  | 747 | 958.509 | 838.51 | 908.8892 | 1576.1887 | 9380 | 1145 | 0.14286 | 3 | 0 | 16.8737 | - | 3.1981 | 3.1981
60 |  | 539 (539/0) |  |  |  | 89.1069 | 59.6371 | 72.065 | 109.6753 | 20305 | 9 | 0.070501 | 16 | 0 | 4.9954 | - | - | -
61 |  | 539 (539/0) |  |  |  | 89.3524 | 58.4335 | 72.1947 | 62.3117 | 13814 | 9 | 0.03154 | 11 | 0 | 5.1791 | - | - | -
62 |  | 538 (538/0) |  |  |  | 89.9126 | 56.7548 | 72.0285 | 46.1822 | 14059 | 0 | 0.020446 | 11 | 0 | 5.1894 | - | - | -
63 |  | 538 (538/0) |  |  |  | 89.2548 | 59.5547 | 72.5794 | 48.6134 | 7720 | 9 | 0.022305 | 6 | 0 | 5.2591 | - | - | -
64 |  | 538 (538/0) |  |  |  | 77.6475 | -71.2046 | 18.964 | 2037.7156 | 2340 | 1929 | 1 | 1 | 1 | 17.7726 | - | - | -
65 |  | 538 (538/0) |  |  |  | 185.9379 | -118.3865 | 34.4416 | 20.2862 | 79 | 6 | 0 | 0 | 0 | 67.8202 | - | - | -
66 |  | 538 (538/0) |  |  |  | 88.7622 | -27.9859 | 21.6844 | 20.8364 | 69 | 9 | 0 | 0 | 0 | 16.8428 | - | - | -
67 |  | 538 (538/0) |  |  |  | 192.9565 | -54.0677 | 27.8627 | 20.9442 | 79 | 8 | 0 | 0 | 0 | 29.0358 | - | - | -
68 |  | 534 (534/0) |  |  |  | 199.568 | 146.0467 | 170.3562 | 192.03 | 13850 | 9 | 0.13296 | 11 | 0 | 10.927 | - | - | -
69 |  | 531 (531/0) |  |  |  | 199.543 | 145.8546 | 171.993 | 302.5895 | 14063 | 6 | 0.21846 | 11 | 0 | 11.1356 | - | - | -
70 |  | 529 (529/0) |  |  |  | 199.2355 | 145.5604 | 170.8543 | 274.3365 | 14202 | 9 | 0.19849 | 11 | 0 | 11.254 | - | - | -
71 |  | 524 (524/0) |  |  |  | 199.6841 | 145.4129 | 172.0833 | 706.5706 | 20726 | 8 | 0.53244 | 16 | 0 | 10.7763 | - | - | -
72 | D109 – 540nm Brightness (20mA) Measure diff with V_ref (temp) | 562 (562/0) | mV |  |  | 833.787 | 745.754 | 791.0584 | 42.1512 | 138 | 19 | 0 | 0 | 0 | 15.4486 | - | - | -
73 | D109 – 630nm Brightness (20mA) Measure diff with V_ref | 559 (559/0) | mV |  | 542 | 849.548 | 654.543 | 794.3479 | 3777.9141 | 11911 | 1232 | 1.0948 | 4 | 0 | 26.1408 | - | 3.2178 | 3.2178
74 | D109 – 660nm Brightness (20mA) Measure diff with V_ref | 558 (558/0) | mV |  | 711 | 1036.182 | 895.885 | 964.509 | 3989.1057 | 12029 | 1229 | 1.1344 | 4 | 0 | 23.0094 | - | 3.6725 | 3.6725
75 | D109 – 940nm Brightness (20mA) Measure diff with V_ref | 556 (556/0) | mV |  | 639 | 997.69 | 858.896 | 932.1818 | 3744.2356 | 15350 | 1234 | 1.0755 | 4 | 0 | 23.395 | - | 4.1773 | 4.1773
76 | D110 – 540nm Brightness (20mA) Measure diff with V_ref | 554 (554/0) | mV |  | 693 | 821.465 | 744.708 | 782.4327 | 3837.1209 | 11726 | 1249 | 1.1173 | 4 | 0 | 13.9057 | - | 2.1438 | 2.1438
77 | D110 – 630nm Brightness (20mA) Measure diff with V_ref | 551 (551/0) | mV |  | 542 | 833.137 | 641.906 | 789.5762 | 3751.8693 | 11340 | 1240 | 1.0835 | 4 | 0 | 23.6444 | - | 3.4903 | 3.4903
78 | D110 – 660nm Brightness (20mA) Measure diff with V_ref | 550 (550/0) | mV |  | 711 | 1020.065 | 919.652 | 960.5731 | 3802.7745 | 11233 | 1273 | 1.1055 | 4 | 0 | 20.3384 | - | 4.0903 | 4.0903
79 |  | 548 (548/0) |  |  |  | 999.336 | 849.215 | 929.1455 | 3755.7628 | 11316 | 1232 | 1.0858 | 4 | 0 | 23.0851 | - | - | -
80 |  | 547 (547/0) |  |  |  | 237.462 | 197.805 | 218.0139 | 3781.6764 | 12106 | 1235 | 1.0951 | 4 | 0 | 6.0387 | - | - | -
81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | 562 (562/0) | mV |  | 693 | 833.787 | 745.754 | 791.0584 | 3740.5356 | 11609 | 1233 | 1.0783 | 4 | 0 | 15.4486 | - | 2.1158 | 2.1158
82 | D110 – 540nm Brightness (20mA) Measure diff with V_ref (temp) | 554 (554/0) | mV |  | 100 | 821.465 | 744.708 | 782.4327 | 41.9477 | 100 | 19 | 0 | 0 | 0 | 13.9057 | - | 16.3585 | 16.3585
83 | Inner green LED brightness AVE | 554 (554/0) | mV |  |  | 818.9425 | 753.629 | 786.7285 | 1183.3953 | 1459 | 1121 | 0 | 0 | 0 | 11.9303 | - | - | -
84 |  | 547 (547/0) |  |  |  | 21133.44 | 18381.74 | 19956.1173 | 38.6362 | 220 | 19 | 0 | 0 | 0 | 451.9708 | - | - | -
85 |  | 540 (540/0) |  |  |  | 20826.92 | 18642.05 | 19954.1198 | 71.0259 | 200 | 39 | 0 | 0 | 0 | 344.3693 | - | - | -
86 | UART連線¦UART Connection | 824 (824/0) |  |  |  | 1 | 1 | 1 | 540.5121 | 719 | 480 | 0 | 0 | 0 | 0 | - | - | -
87 | Is ESN check in | 0 (0/0) |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | -
88 | AFE Power Consumption data | 575 (575/0) |  |  | 1000 | 3380.702 | 2297.389 | 2924.6732 | 2051.633 | 10325 | 1939 | 1.0261 | 9 | 1 | 275.1269 | - | 2.3319 | 2.3319
89 | AFE Power Consumption Test sampling rate to 128Hz | 572 (572/0) |  | 128 | 128 | 128 | 128 | 128 | 279.9003 | 620 | 138 | 0 | 0 | 0 | 0 | - | - | -
90 | D105 (540nm) Brightness and LED Current (20mA) | 572 (572/0) | mA | 22 | 18 | 21.229 | 20.335 | 20.8572 | 1138.4371 | 1339 | 1084 | 0 | 0 | 0 | 0.20053 | -0.42858 | 3.3246 | 1.8997
91 | D105 (540nm) Brightness and LED | 572 (572/0) |  | 1 | 1 | 1 | 1 | 1 | 605.9266 | 748 | 555 | 0 | 0 | 0 | 0 | - | - | -
92 | D106 – 540nm Brightness (20mA) | 572 (572/0) |  | 1 | 1 | 1 | 1 | 1 | 609.8829 | 723 | 561 | 0 | 0 | 0 | 0 | - | - | -
94 |  | 547 (547/0) |  |  |  | 1 | 1 | 1 | 172.2633 | 685 | 139 | 0 | 0 | 0 | 0 | - | - | -
95 | D109 – 940nm Brightness (20mA) | 557 (557/0) |  | 1 | 1 | 1 | 1 | 1 | 609.5332 | 963 | 566 | 0 | 0 | 0 | 0 | - | - | -
96 | D109 – 660nm Brightness (20mA) | 559 (559/0) |  | 1 | 1 | 1 | 1 | 1 | 606.4061 | 764 | 561 | 0 | 0 | 0 | 0 | - | - | -
97 | D109 – 630nm Brightness (20mA) | 561 (561/0) |  | 1 | 1 | 1 | 1 | 1 | 609.7255 | 754 | 564 | 0 | 0 | 0 | 0 | - | - | -
98 | D109 – 540nm Brightness (20mA) | 566 (566/0) |  | 1 | 1 | 1 | 1 | 1 | 607.9258 | 753 | 567 | 0 | 0 | 0 | 0 | - | - | -
99 | D108 – 540nm Brightness (20mA) | 570 (570/0) |  | 1 | 1 | 1 | 1 | 1 | 608.3842 | 785 | 559 | 0 | 0 | 0 | 0 | - | - | -
100 | D107 – 540nm Brightness (20mA) | 570 (570/0) |  | 1 | 1 | 1 | 1 | 1 | 606.3544 | 837 | 559 | 0 | 0 | 0 | 0 | - | - | -
101 |  | 548 (548/0) |  |  |  | 1 | 1 | 1 | 606.5237 | 764 | 562 | 0 | 0 | 0 | 0 | - | - | -
102 | D110 – 940nm Brightness (20mA) | 550 (550/0) |  | 1 | 1 | 1 | 1 | 1 | 608.7164 | 910 | 566 | 0 | 0 | 0 | 0 | - | - | -
103 | D110 – 660nm Brightness (20mA) | 551 (551/0) |  | 1 | 1 | 1 | 1 | 1 | 608.7441 | 859 | 562 | 0 | 0 | 0 | 0 | - | - | -
104 | D110 – 630nm Brightness (20mA) | 553 (553/0) |  | 1 | 1 | 1 | 1 | 1 | 607.4376 | 753 | 551 | 0 | 0 | 0 | 0 | - | - | -
105 | D110 – 540nm Brightness (20mA) | 556 (556/0) |  | 1 | 1 | 1 | 1 | 1 | 608.0234 | 863 | 562 | 0 | 0 | 0 | 0 | - | - | -
106 |  | 524 (524/0) |  |  |  | 3500.558 | 2488.026 | 3043.4028 | 20.9256 | 120 | 0 | 0 | 0 | 0 | 221.8848 | - | - | -
107 |  | 524 (524/0) |  |  |  | 2605.2 | 1896.571 | 2377.3821 | 20.7595 | 79 | 9 | 0 | 0 | 0 | 96.9422 | - | - | -
108 |  | 524 (524/0) |  |  |  | 3436.935 | 2792.59 | 3093.2305 | 21.1775 | 90 | 9 | 0 | 0 | 0 | 108.9514 | - | - | -
109 |  | 524 (524/0) |  |  |  | 2727.724 | 1942.706 | 2528.0633 | 21.4733 | 99 | 9 | 0 | 0 | 0 | 106.894 | - | - | -
110 |  | 524 (524/0) |  |  |  | 3396.913 | 2751.178 | 3097.1645 | 21.3454 | 87 | 5 | 0 | 0 | 0 | 113.4099 | - | - | -
111 |  | 524 (524/0) |  |  |  | 3380.702 | 2297.389 | 2932.1789 | 20.8378 | 80 | 9 | 0 | 0 | 0 | 272.5411 | - | - | -
112 |  | 523 (523/0) |  |  |  | 14695.34 | 10144.15 | 12813.2392 | 2020.3767 | 7634 | 1920 | 1.0096 | 6 | 1 | 1176.5484 | - | - | -
113 |  | 523 (523/0) |  |  |  | 15114.77 | 10936.14 | 13299.2047 | 20.6023 | 80 | 9 | 0 | 0 | 0 | 961.3823 | - | - | -
114 |  | 523 (523/0) |  |  |  | 15438.12 | 11278.22 | 14057.4453 | 20.5545 | 100 | 5 | 0 | 0 | 0 | 543.0816 | - | - | -
115 |  | 523 (523/0) |  |  |  | 20817.11 | 16931.8 | 18741.9181 | 21.6692 | 139 | 9 | 0 | 0 | 0 | 649.0908 | - | - | -
116 |  | 523 (523/0) |  |  |  | 16108.62 | 11611.96 | 14935.4021 | 20.9446 | 79 | 9 | 0 | 0 | 0 | 604.3221 | - | - | -
117 |  | 523 (523/0) |  |  |  | 0.75633 | 0.70459 | 0.72152 | 133.1166 | 520 | 86 | 0 | 0 | 0 | 0.011714 | - | - | -
118 |  | 523 (523/0) |  |  |  | 0.77525 | 0.71353 | 0.73963 | 133.0172 | 469 | 89 | 0 | 0 | 0 | 0.013621 | - | - | -
119 |  | 523 (523/0) |  |  |  | 0.75369 | 0.70367 | 0.7212 | 135.0402 | 430 | 79 | 0 | 0 | 0 | 0.011646 | - | - | -
120 |  | 523 (523/0) |  |  |  | 0.7748 | 0.71708 | 0.73893 | 137.0249 | 437 | 79 | 0 | 0 | 0 | 0.013715 | - | - | -
121 |  | 523 (523/0) |  |  |  | 20540.3 | 16711.71 | 18758.3946 | 20.2677 | 79 | 8 | 0 | 0 | 0 | 678.1906 | - | - | -

## ECG Test Item
Code | Name | Count | Unit | USL | LSL | MaxValue | MinValue | AvgValue | AvgSpare | MaxSpare | MinSpare | AvgRetries | MaxRetries | MinRetries | Stddev | Ck | Cp | Cpk
1 | Final Index 32766 | 607 (607/0) |  | 32766 | 32766 | 32766 | 32766 | 32766 | 44.2026 | 99 | 35 | 0 | 0 | 0 | 0 | - | - | -
2 | UART連線com1¦UART Connection | 607 (607/0) |  |  |  | 1 | 1 | 1 | 521.3937 | 719 | 480 | 0 | 0 | 0 | 0 | - | - | -
3 | UART連線com2¦UART Connection | 607 (607/0) |  |  |  | 1 | 1 | 1 | 3733.7677 | 10360 | 3475 | 0.0098847 | 3 | 0 | 0 | - | - | -
4 | initialize_new | 607 (607/0) |  |  |  | 0 | 0 | 0 | 1.2488 | 20 | 0 | 0 | 0 | 0 | 0 | - | - | -
5 | FID | 607 (607/0) |  |  |  | 255 | 2 | 60.2784 | 1197.7232 | 1280 | 1125 | 0 | 0 | 0 | 105.6276 | - | - | -
6 | SW version | 607 (607/0) |  |  | 2.17 | 7.95 | 7.95 | 7.95 | 552.6557 | 620 | 517 | 0 | 0 | 0 | 9.0594e-14 | - | 2.1267e+13 | 2.1267e+13
7 | USB Connection (1) | 606 (606/0) |  |  |  | 1 | 1 | 1 | 19029.099 | 25240 | 17181 | 0.016502 | 5 | 0 | 0 | - | - | -
8 | Upload ESN | 0 (0/0) |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | -
9 | Is ESN check in | 548 (548/0) |  |  |  | 0 | 0 | 0 | 200.8449 | 520 | 178 | 0 | 0 | 0 | 0 | - | - | -
10 | Start WF_geneator 10Hz 500mV | 548 (548/0) |  |  |  | 0 | 0 | 0 | 670.0201 | 719 | 629 | 0 | 0 | 0 | 0 | - | - | -
11 | WF_geneator Connection (1) | 548 (548/0) |  |  |  | 1 | 1 | 1 | 2416.615 | 2639 | 2282 | 0 | 0 | 0 | 0 | - | - | -
12 | USB Connection (2) | 548 (548/0) |  |  |  | 1 | 1 | 1 | 513.3485 | 550 | 19 | 0 | 0 | 0 | 0 | - | - | -
13 | OFF com C314 (1) | 548 (548/0) |  |  |  | 0 | 0 | 0 | 1637.4544 | 1719 | 1530 | 0 | 0 | 0 | 0 | - | - | -
14 | ECG RLD 10Hz 500mV Test | 543 (543/0) | mV | 500 | 350 | 488 | 420 | 435.0976 | 1430.9982 | 1522 | 1399 | 1 | 1 | 1 | 14.889 | -0.13463 | 1.6791 | 1.453
15 | Check 1907IOP | 543 (543/0) | mV | 6 | 6 | 6 | 6 | 6 | 565.8582 | 679 | 529 | 0 | 0 | 0 | 0 | - | - | -
16 | WF_geneator Connection (2) | 543 (543/0) |  |  |  | 1 | 1 | 1 | 1823.8895 | 2869 | 1726 | 0.016575 | 1 | 0 | 0 | - | - | -
17 | ON com C314 (1) | 543 (543/0) |  |  |  | 0 | 0 | 0 | 1638.0626 | 1718 | 1531 | 0 | 0 | 0 | 0 | - | - | -
18 | OFF com C314 (2) | 543 (543/0) |  |  |  | 0 | 0 | 0 | 2945.9576 | 3144 | 2751 | 0 | 0 | 0 | 0 | - | - | -
19 | Start WF_geneator 5Hz 2mV | 543 (543/0) |  |  |  | 0 | 0 | 0 | 1105.3131 | 1150 | 1036 | 0 | 0 | 0 | 0 | - | - | -
20 | ECG 2mv 5Hz Test | 536 (536/0) | mv | 2.2 | 1.9 | 2.198 | 1.9005 | 2.0939 | 6009.6362 | 35787 | 4689 | 0.1959 | 5 | 0 | 0.033842 | -0.29283 | 1.4774 | 1.0448
21 | USB Connection (3) | 543 (543/0) |  |  |  | 1 | 1 | 1 | 494.3738 | 550 | 13 | 0 | 0 | 0 | 0 | - | - | -
22 | ON com C314 (3) | 536 (536/0) |  |  |  | 0 | 0 | 0 | 3809.3955 | 3940 | 3565 | 0 | 0 | 0 | 0 | - | - | -
23 | WF_geneator Connection (3) | 536 (536/0) |  |  |  | 1 | 1 | 1 | 1755.6978 | 1919 | 1677 | 0 | 0 | 0 | 0 | - | - | -
24 | Start WF_geneator 1Hz 2mV | 536 (536/0) |  |  |  | 0 | 0 | 0 | 1105.2276 | 1150 | 1037 | 0 | 0 | 0 | 0 | - | - | -
25 | OFF com C314 (2) | 536 (536/0) |  |  |  | 0 | 0 | 0 | 2945.9086 | 3049 | 2757 | 0 | 0 | 0 | 0 | - | - | -
26 | USB Connection (4) | 536 (536/0) |  |  |  | 1 | 1 | 1 | 493.9832 | 559 | 10 | 0 | 0 | 0 | 0 | - | - | -
27 | ECG 2mv 1Hz Test | 535 (535/0) | mv | 2.2 | 1.65 | 2.0504 | 1.704 | 1.8482 | 4603.7888 | 10599 | 4489 | 0.0037383 | 1 | 0 | 0.044587 | 0.27913 | 2.0559 | 1.482
28 | ON com C314 (4) | 535 (535/0) |  |  |  | 0 | 0 | 0 | 2719.0262 | 2849 | 2558 | 0 | 0 | 0 | 0 | - | - | -
29 | WF_geneator Connection (4) | 535 (535/0) |  |  |  | 1 | 1 | 1 | 1755.7364 | 1929 | 1671 | 0 | 0 | 0 | 0 | - | - | -
30 | Motor Resonant Frequency | 551 (551/0) | Hz | 210 | 185 | 199 | 185 | 189.8076 | 1143.3612 | 28646 | 1018 | 0.010889 | 6 | 0 | 2.7908 | 0.61539 | 1.493 | 0.57422
32 | Input ESN | 607 (607/0) |  |  |  | 0 | 0 | 0 | 15072.2175 | 2132757 | 3447 | 0 | 0 | 0 | 0 | - | - | -
54 | Check Fixture ID from registry | 0 (0/0) |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | -
55 | Start WF_geneator 2Hz 2mV | 535 (535/0) |  |  |  | 0 | 0 | 0 | 1104.757 | 1170 | 1036 | 0 | 0 | 0 | 0 | - | - | -
56 | Exclude Golden ESN | 0 (0/0) |  | 0 | 0 | - | - | - | - | - | - | - | - | - | - | - | - | -
57 | OFF com C314 (4) | 535 (535/0) |  |  |  | 0 | 0 | 0 | 2946.8991 | 3163 | 2758 | 0 | 0 | 0 | 0 | - | - | -
58 | USB Connection (5) | 535 (535/0) |  |  |  | 1 | 1 | 1 | 498.4561 | 540 | 19 | 0 | 0 | 0 | 0 | - | - | -
59 | ECG 2mv 2Hz Test | 533 (533/0) | mv | 2.2 | 1.85 | 2.1919 | 1.9669 | 2.0528 | 4659.9343 | 10615 | 4489 | 0.013133 | 1 | 0 | 0.04696 | -0.15911 | 1.2422 | 1.0446
60 | ON com C314 (5) | 533 (533/0) |  |  |  | 0 | 0 | 0 | 2720.6754 | 2869 | 2547 | 0 | 0 | 0 | 0 | - | - | -
61 | WF_geneator Connection (5) | 533 (533/0) |  |  |  | 1 | 1 | 1 | 1754.06 | 1929 | 1664 | 0 | 0 | 0 | 0 | - | - | -
62 | Start WF_geneator 40Hz 2mV | 533 (533/0) |  |  |  | 0 | 0 | 0 | 1105.0319 | 1190 | 1039 | 0 | 0 | 0 | 0 | - | - | -
63 | OFF com C314 (5) | 533 (533/0) |  |  |  | 0 | 0 | 0 | 2947.0432 | 3100 | 2752 | 0 | 0 | 0 | 0 | - | - | -
64 | USB Connection (6) | 533 (533/0) |  |  |  | 1 | 1 | 1 | 505.3846 | 550 | 19 | 0 | 0 | 0 | 0 | - | - | -
65 | ECG 2mv 40Hz Test | 533 (533/0) | mv | 2.2 | 1.85 | 2.1812 | 1.9772 | 2.0536 | 4626.3265 | 10622 | 4489 | 0.0075047 | 1 | 0 | 0.054037 | -0.16364 | 1.0795 | 0.90285
66 | ON com C314 (5) | 533 (533/0) |  |  |  | 0 | 0 | 0 | 2720.3021 | 2860 | 2545 | 0 | 0 | 0 | 0 | - | - | -
67 | WF_geneator Connection (6) | 533 (533/0) |  |  |  | 1 | 1 | 1 | 1535.9906 | 1710 | 1450 | 0 | 0 | 0 | 0 | - | - | -
68 | Start WF_geneator 0.4Hz 2mV | 533 (533/0) |  |  |  | 0 | 0 | 0 | 1105.5366 | 1159 | 1037 | 0 | 0 | 0 | 0 | - | - | -
69 | OFF com C314 (6) | 533 (533/0) |  |  |  | 0 | 0 | 0 | 2945.8874 | 3010 | 2753 | 0 | 0 | 0 | 0 | - | - | -
70 | USB Connection (7) | 533 (533/0) |  |  |  | 1 | 1 | 1 | 505.2645 | 550 | 12 | 0 | 0 | 0 | 0 | - | - | -
71 | ECG 2mV 0.4 Hz Test | 533 (533/0) | mv | 1.2 | 0.85 | 1.1715 | 0.95901 | 1.053 | 4637.1107 | 10639 | 4488 | 0.0093809 | 1 | 0 | 0.034018 | -0.16023 | 1.7148 | 1.44
72 | == OFF ECG Test == | 533 (533/0) |  |  |  | 0 | 0 | 0 | 3.4578 | 20 | 0 | 0 | 0 | 0 | 0 | - | - | -

## Critical Part Record (hidden)
Part 5 : Detail of Critical Part
Critical part definition
1. EE,ME,DE provide critical parts list
N/A |  |  |  |  | Factory Prototype Run | Factory Prototype Run | Factory Prototype Run | Factory Prototype Run | Factory Prototype Run | Pilot Run | Pilot Run | Pilot Run | Pilot Run | Pilot Run
Part | GPN | Description | Type | Comments (Detail definition of part) | Failed Rate | Tracking in Prweb(Y/N) | ENG Part (Y/N) | Vender | Vender's Date Code | Failed Rate | Tracking in Prweb(Y/N) | ENG Part (Y/N) | Vender | Vender's Date Code

## Parameter and Fixture tracking (hidden)
Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking | Part 6: Parameter and Fixture tracking
PFMEA NO. | Process | Station | Description | FFP | FFP | FFP | FFP | FFP | Pilot Run | Pilot Run | Pilot Run | Pilot Run | Pilot Run | MP | MP | MP | MP | MP
PFMEA NO. | Process | Station | Description | Change Reason | Parameter | Fixture Name | Fixture ID | Critical Manner | Change Reason | Parameter | Fixture Name | Fixture ID | Critical Manner | Change Reason | Parameter | Fixture Name | Fixture ID | Critical Manner

## Raw data
JobNO | ItemNameType | StationID | SerialNumber | tDateTime | Result | FixtureID2 | failitem | TestType | TestType2 | Barometer Temperature | AFE Power Consumption data
16045256 | 17859 | 4 | 2905E04594 | 2022-09-23T01:04:20.657+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04942 | 2022-09-23T01:02:17.84+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04690 | 2022-09-23T00:48:55.513+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04661 | 2022-09-23T00:45:03.967+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04451 | 2022-09-23T00:43:43.477+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04451 | 2022-09-23T00:42:10.31+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04958 | 2022-09-23T00:41:45.423+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04960 | 2022-09-23T00:39:59.07+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03852 | 2022-09-23T00:38:59.053+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03579 | 2022-09-23T00:19:13.427+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04014 | 2022-09-22T11:05:32.157+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03579 | 2022-09-22T11:05:14.653+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E03579 | 2022-09-22T10:41:42.72+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E04773 | 2022-09-22T10:40:01.67+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04014 | 2022-09-22T10:38:05.107+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E04095 | 2022-09-22T10:33:46.34+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03649 | 2022-09-22T10:30:09.423+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03982 | 2022-09-22T10:26:40.547+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04119 | 2022-09-22T10:25:32.56+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E05296 | 2022-09-22T10:19:47.583+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E04095 | 2022-09-22T10:18:53.617+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E03649 | 2022-09-22T10:14:49.32+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03579 | 2022-09-22T10:09:50.607+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04773 | 2022-09-22T10:04:01.727+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E03744 | 2022-09-22T10:02:46.183+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E04014 | 2022-09-22T10:02:35.857+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03762 | 2022-09-22T08:56:09.993+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03762 | 2022-09-22T08:50:27.173+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04471 | 2022-09-22T08:41:11.263+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03763 | 2022-09-22T08:39:50.517+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04471 | 2022-09-22T08:37:18.39+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E03763 | 2022-09-22T08:34:48.223+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04107 | 2022-09-22T08:29:51.107+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E03971 | 2022-09-22T08:27:42.817+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03971 | 2022-09-22T08:26:14.973+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E05247 | 2022-09-22T08:22:48.5+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E05247 | 2022-09-22T08:21:18.947+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03773 | 2022-09-22T08:00:50.47+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E03773 | 2022-09-22T07:59:37.237+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03940 | 2022-09-22T07:59:15.9+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04677 | 2022-09-22T07:41:02.073+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E05171 | 2022-09-22T07:36:32.657+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04011 | 2022-09-22T07:34:02.027+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E05171 | 2022-09-22T07:33:40.337+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03746 | 2022-09-22T07:23:45.747+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E03746 | 2022-09-22T07:19:30.867+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04653 | 2022-09-22T07:18:25.817+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03926 | 2022-09-22T07:18:07.33+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04469 | 2022-09-22T07:16:08.937+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03926 | 2022-09-22T07:15:01.797+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04690 | 2022-09-22T06:45:43.173+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04451 | 2022-09-22T06:43:19.22+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04960 | 2022-09-22T06:43:06.647+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03991 | 2022-09-22T06:42:35.157+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03852 | 2022-09-22T06:37:55.967+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03991 | 2022-09-22T06:34:07.393+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03705 | 2022-09-22T06:25:11.873+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04513 | 2022-09-22T06:09:56.99+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E05086 | 2022-09-22T06:04:02.533+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04034 | 2022-09-22T06:02:28.93+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E05086 | 2022-09-22T05:59:05.433+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04034 | 2022-09-22T05:57:48.7+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E03924 | 2022-09-22T05:52:13.133+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04547 | 2022-09-22T05:01:44.827+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04946 | 2022-09-22T04:28:12.193+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04144 | 2022-09-22T04:23:47.287+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04000 | 2022-09-22T04:06:35.537+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E05124 | 2022-09-22T03:31:22.767+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E05285 | 2022-09-22T03:28:39.997+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E03852 | 2022-09-22T02:40:52.083+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04982 | 2022-09-22T02:40:02.017+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E03852 | 2022-09-22T02:37:04.237+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E05119 | 2022-09-22T02:28:36.257+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04742 | 2022-09-22T02:27:11.273+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E03134 | 2022-09-22T02:25:29.783+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04451 | 2022-09-22T02:18:18.743+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04451 | 2022-09-22T02:10:52.267+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E04321 | 2022-09-22T01:50:30.853+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 2 | 2905E04958 | 2022-09-22T01:40:45.493+00:00 | FALSE | 3 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E01152 | 2022-09-22T01:39:00.773+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04000 | 2022-09-22T01:29:04.88+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E03777 | 2022-09-22T01:27:38.237+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04507 | 2022-09-22T01:23:17.73+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04459 | 2022-09-22T01:21:55.83+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04520 | 2022-09-22T01:19:33.24+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04575 | 2022-09-22T01:18:22.677+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04593 | 2022-09-22T01:16:29.9+00:00 | FALSE | 4 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E04144 | 2022-09-22T01:15:44.81+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E04646 | 2022-09-22T00:54:24.807+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 3 | 2905E04942 | 2022-09-22T00:34:31.327+00:00 | FALSE | 2 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 1 | 2905E00996 | 2022-09-21T10:18:51.62+00:00 | FALSE | 1 | 41 | . | 152 | -999 | 0
16045256 | 17859 | 4 | 2905E04594 | 2022-09-23T00:19:42.68+00:00 | FALSE | 4 | 41 | . | 152 | -270.8047 | 0
16045256 | 17859 | 4 | 2905E04578 | 2022-09-22T09:44:29.797+00:00 | FALSE |  | 3 | . | 152 | 0 | 0
16045256 | 17859 | 2 | 2905E05048 | 2022-09-22T09:44:17.847+00:00 | FALSE |  | 3 | . | 152 | 0 | 0
16045256 | 17859 | 3 | 2905E01253 | 2022-09-21T09:59:20.477+00:00 | FALSE | 2 | 4 | . | 152 | 0 | 0
16045256 | 17859 | 1 | 2905E01013 | 2022-09-21T09:59:20.397+00:00 | FALSE | 1 | 4 | . | 152 | 0 | 0
16045256 | 17859 | 1 | 2905E01013 | 2022-09-21T09:54:50.267+00:00 | FALSE |  | 3 | . | 152 | 0 | 0
16045256 | 17859 | 1 | 2905E04035 | 2022-09-21T07:07:35.87+00:00 | FALSE | 3 | 4 | . | 152 | 0 | 0
16045256 | 17859 | 2 | 2905E04685 | 2022-09-22T01:07:04.207+00:00 | TRUE | 3 | 0 | . | 152 | 21.2109 | 3226.044
16045256 | 17859 | 4 | 2905E04546 | 2022-09-22T01:07:18.333+00:00 | TRUE | 4 | 0 | . | 152 | 21.2188 | 2497.333
16045256 | 17859 | 2 | 2905E04894 | 2022-09-22T01:09:52.933+00:00 | TRUE | 3 | 0 | . | 152 | 21.4531 | 3175.599
16045256 | 17859 | 2 | 2905E04354 | 2022-09-22T01:14:21.807+00:00 | FALSE | 3 | 88 | . | 152 | 21.5312 | -999
16045256 | 17859 | 4 | 2905E04340 | 2022-09-22T01:12:51.667+00:00 | TRUE | 4 | 0 | . | 152 | 21.6953 | 2450.692
16045256 | 17859 | 2 | 2905E04361 | 2022-09-22T01:12:43.317+00:00 | TRUE | 3 | 0 | . | 152 | 21.7109 | 3334.665
16045256 | 17859 | 2 | 2905E04545 | 2022-09-22T01:17:29.293+00:00 | TRUE | 3 | 0 | . | 152 | 21.7812 | 3138.914
16045256 | 17859 | 3 | 2905E03874 | 2022-09-22T00:12:24.447+00:00 | TRUE | 2 | 0 | . | 152 | 21.8672 | 3064.802
16045256 | 17859 | 4 | 2905E04456 | 2022-09-22T01:10:06.017+00:00 | TRUE | 4 | 0 | . | 152 | 21.875 | 2448.197
16045256 | 17859 | 4 | 2905E04481 | 2022-09-22T01:15:27.77+00:00 | TRUE | 4 | 0 | . | 152 | 21.9375 | 2500.965
16045256 | 17859 | 1 | 2905E04661 | 2022-09-22T00:12:41.1+00:00 | FALSE | 1 | 62 | . | 152 | 21.9766 | 3004.249
16045256 | 17859 | 2 | 2905E04762 | 2022-09-22T01:20:05.893+00:00 | TRUE | 3 | 0 | . | 152 | 21.9766 | 3220.772
16045256 | 17859 | 3 | 2905E03979 | 2022-09-22T00:16:26.977+00:00 | TRUE | 2 | 0 | . | 152 | 21.9844 | 3079.837
16045256 | 17859 | 1 | 2905E04182 | 2022-09-22T00:16:29.377+00:00 | TRUE | 1 | 0 | . | 152 | 22.0391 | 2960.081
16045256 | 17859 | 4 | 2905E04014 | 2022-09-23T00:14:17.867+00:00 | TRUE | 4 | 0 | . | 152 | 22.0938 | 2493.648
16045256 | 17859 | 2 | 2905E04104 | 2022-09-22T01:22:49.597+00:00 | TRUE | 3 | 0 | . | 152 | 22.1094 | 3183.849
16045256 | 17859 | 1 | 2905E03940 | 2022-09-23T00:14:24.587+00:00 | TRUE | 1 | 0 | . | 152 | 22.1875 | 3156.861
16045256 | 17859 | 3 | 2905E03846 | 2022-09-22T00:21:04.11+00:00 | TRUE | 2 | 0 | . | 152 | 22.25 | 3111.772
16045256 | 17859 | 1 | 2905E03272 | 2022-09-22T00:23:42.413+00:00 | TRUE | 1 | 0 | . | 152 | 22.2969 | 3003.308
16045256 | 17859 | 3 | 2905E05129 | 2022-09-22T00:27:35.917+00:00 | FALSE | 2 | 88 | . | 152 | 22.3594 | -999
16045256 | 17859 | 2 | 2905E04547 | 2022-09-22T01:27:46.193+00:00 | FALSE | 3 | 25 | . | 152 | 22.3594 | 3235.393
16045256 | 17859 | 1 | 2905E05247 | 2022-09-23T00:18:06.51+00:00 | TRUE | 1 | 0 | . | 152 | 22.3906 | 3115.872
16045256 | 17859 | 3 | 2905E04692 | 2022-09-22T00:23:57.577+00:00 | TRUE | 2 | 0 | . | 152 | 22.3906 | 3222.113
16045256 | 17859 | 1 | 2905E04737 | 2022-09-22T00:29:58.153+00:00 | FALSE | 1 | 88 | . | 152 | 22.4141 | -999
16045256 | 17859 | 3 | 2905E05329 | 2022-09-22T00:30:08.077+00:00 | TRUE | 2 | 0 | . | 152 | 22.4297 | 2967.706
16045256 | 17859 | 4 | 2905E03773 | 2022-09-23T00:17:39.287+00:00 | TRUE | 4 | 0 | . | 152 | 22.4688 | 2544.12
16045256 | 17859 | 4 | 2905E04632 | 2022-09-23T00:27:21.523+00:00 | TRUE | 4 | 0 | . | 152 | 22.4766 | 2535.192
16045256 | 17859 | 1 | 2905E04674 | 2022-09-22T00:28:06.227+00:00 | TRUE | 1 | 0 | . | 152 | 22.4844 | 3006.205
16045256 | 17859 | 1 | 2905E04811 | 2022-09-23T00:29:39.417+00:00 | TRUE | 1 | 0 | . | 152 | 22.5 | 2947.172
16045256 | 17859 | 1 | 2905E03473 | 2022-09-22T00:20:47.38+00:00 | TRUE | 1 | 0 | . | 152 | 22.5 | 3024.149
16045256 | 17859 | 2 | 2905E04240 | 2022-09-22T01:26:02.59+00:00 | TRUE | 3 | 0 | . | 152 | 22.5156 | 3144.25
16045256 | 17859 | 4 | 2905E04543 | 2022-09-23T00:30:32.993+00:00 | TRUE | 4 | 0 | . | 152 | 22.5625 | 2461.743
16045256 | 17859 | 1 | 2905E04572 | 2022-09-22T00:32:28.62+00:00 | TRUE | 1 | 0 | . | 152 | 22.5625 | 2947.951
16045256 | 17859 | 1 | 2905E04370 | 2022-09-22T00:35:34.123+00:00 | TRUE | 1 | 0 | . | 152 | 22.5938 | 3045.053
16045256 | 17859 | 2 | 2905E04012 | 2022-09-22T01:33:16.347+00:00 | FALSE | 3 | 68 | . | 152 | 22.5938 | 3120.501
16045256 | 17859 | 3 | 2905E04560 | 2022-09-22T00:32:55.36+00:00 | TRUE | 2 | 0 | . | 152 | 22.6562 | 3051.374
16045256 | 17859 | 1 | 2905E04842 | 2022-09-23T00:25:51.813+00:00 | FALSE | 1 | 88 | . | 152 | 22.6797 | -999
16045256 | 17859 | 3 | 2905E04936 | 2022-09-22T00:37:09.41+00:00 | FALSE | 2 | 88 | . | 152 | 22.7188 | -999
16045256 | 17859 | 1 | 2905E04564 | 2022-09-23T00:30:59.59+00:00 | FALSE | 1 | 88 | . | 152 | 22.7422 | -999
16045256 | 17859 | 1 | 2905E03472 | 2022-09-22T00:38:07.787+00:00 | TRUE | 1 | 0 | . | 152 | 22.7422 | 2979.942
16045256 | 17859 | 4 | 2905E04782 | 2022-09-23T00:35:36.377+00:00 | TRUE | 4 | 0 | . | 152 | 22.7656 | 2517.43
16045256 | 17859 | 1 | 2905E04594 | 2022-09-23T00:23:21.417+00:00 | FALSE | 1 | 69 | . | 152 | 22.7656 | 3024.164
16045256 | 17859 | 4 | 2905E03938 | 2022-09-23T00:33:05.08+00:00 | TRUE | 4 | 0 | . | 152 | 22.7812 | 2501.983
16045256 | 17859 | 3 | 2905E03962 | 2022-09-22T01:06:57.753+00:00 | TRUE | 2 | 0 | . | 152 | 22.7891 | 3000.757
16045256 | 17859 | 1 | 2905E03271 | 2022-09-23T00:33:47.587+00:00 | TRUE | 1 | 0 | . | 152 | 22.8203 | 3274.785
16045256 | 17859 | 4 | 2905E01267 | 2022-09-22T01:28:26.527+00:00 | FALSE | 4 | 21 | . | 152 | 22.8438 | 2390.479
16045256 | 17859 | 4 | 2905E04866 | 2022-09-23T00:36:49.577+00:00 | FALSE | 4 | 46 | . | 152 | 22.8516 | 0
16045256 | 17859 | 1 | 2905E04866 | 2022-09-23T00:34:53.837+00:00 | FALSE | 1 | 46 | . | 152 | 22.8516 | 0
16045256 | 17859 | 4 | 2905E03963 | 2022-09-22T01:26:40.2+00:00 | TRUE | 4 | 0 | . | 152 | 22.8516 | 2415.1
16045256 | 17859 | 4 | 2905E04139 | 2022-09-23T00:23:53.827+00:00 | FALSE | 4 | 10 | . | 152 | 22.8516 | 2487.613
16045256 | 17859 | 1 | 2905E03946 | 2022-09-23T00:37:53.497+00:00 | TRUE | 1 | 0 | . | 152 | 22.8594 | 3005.959
16045256 | 17859 | 3 | 2905E03541 | 2022-09-22T00:39:53.493+00:00 | TRUE | 2 | 0 | . | 152 | 22.8672 | 3025.6
16045256 | 17859 | 4 | 2905E03971 | 2022-09-23T00:39:15.66+00:00 | TRUE | 4 | 0 | . | 152 | 22.8828 | 2533.686
16045256 | 17859 | 3 | 2905E04958 | 2022-09-22T00:41:33.883+00:00 | FALSE | 2 | 88 | . | 152 | 22.9844 | -999
16045256 | 17859 | 4 | 2905E01440 | 2022-09-22T01:43:26.907+00:00 | TRUE | 4 | 0 | . | 152 | 22.9844 | 2448.478
16045256 | 17859 | 4 | 2905E04050 | 2022-09-22T01:33:27.217+00:00 | TRUE | 4 | 0 | . | 152 | 23 | 2415.493
16045256 | 17859 | 3 | 2905E05180 | 2022-09-22T01:09:31.497+00:00 | TRUE | 2 | 0 | . | 152 | 23 | 2992.084
16045256 | 17859 | 1 | 2905E05027 | 2022-09-22T01:06:45.93+00:00 | TRUE | 1 | 0 | . | 152 | 23.0156 | 2994.477
16045256 | 17859 | 4 | 2905E01265 | 2022-09-22T01:40:39.703+00:00 | TRUE | 4 | 0 | . | 152 | 23.0312 | 2391.887
16045256 | 17859 | 4 | 2905E03705 | 2022-09-23T00:54:01.667+00:00 | FALSE | 4 | 88 | . | 152 | 23.0547 | -999
16045256 | 17859 | 3 | 2905E04839 | 2022-09-22T00:52:52.427+00:00 | FALSE | 2 | 88 | . | 152 | 23.0859 | -999
16045256 | 17859 | 2 | 2905E00942 | 2022-09-22T01:36:01.623+00:00 | TRUE | 3 | 0 | . | 152 | 23.0859 | 3170.99
16045256 | 17859 | 4 | 2905E04842 | 2022-09-23T00:40:21.873+00:00 | FALSE | 4 | 88 | . | 152 | 23.1016 | -999
16045256 | 17859 | 1 | 2905E04673 | 2022-09-22T00:42:54.05+00:00 | TRUE | 1 | 0 | . | 152 | 23.1016 | 2986.456
16045256 | 17859 | 4 | 2905E04737 | 2022-09-22T01:36:34.397+00:00 | TRUE | 4 | 0 | . | 152 | 23.1094 | 2410.291
16045256 | 17859 | 1 | 2905E04012 | 2022-09-23T00:50:49.787+00:00 | TRUE | 1 | 0 | . | 152 | 23.1328 | 2924.135
16045256 | 17859 | 3 | 2905E03143 | 2022-09-22T00:44:33.663+00:00 | TRUE | 2 | 0 | . | 152 | 23.1328 | 2944.297
16045256 | 17859 | 2 | 2905E04488 | 2022-09-22T01:46:57.867+00:00 | TRUE | 3 | 0 | . | 152 | 23.1406 | 3167.716
16045256 | 17859 | 2 | 2905E04241 | 2022-09-22T01:43:55.813+00:00 | TRUE | 3 | 0 | . | 152 | 23.1797 | 3115.436
16045256 | 17859 | 3 | 2905E04850 | 2022-09-22T00:51:36.367+00:00 | FALSE | 2 | 88 | . | 152 | 23.2031 | -999
16045256 | 17859 | 3 | 2905E04566 | 2022-09-22T01:14:36.007+00:00 | TRUE | 2 | 0 | . | 152 | 23.2109 | 2865.403
16045256 | 17859 | 1 | 2905E04942 | 2022-09-23T00:56:45.34+00:00 | FALSE | 1 | 88 | . | 152 | 23.2188 | -999
16045256 | 17859 | 1 | 2905E04690 | 2022-09-23T00:46:46.197+00:00 | FALSE | 1 | 88 | . | 152 | 23.2188 | -999
16045256 | 17859 | 4 | 2905E04958 | 2022-09-23T00:45:16.107+00:00 | TRUE | 4 | 0 | . | 152 | 23.2188 | 2514.215
16045256 | 17859 | 4 | 2905E03881 | 2022-09-22T01:55:15.183+00:00 | TRUE | 4 | 0 | . | 152 | 23.2266 | 2453.217
16045256 | 17859 | 1 | 2905E03886 | 2022-09-22T00:52:40.607+00:00 | TRUE | 1 | 0 | . | 152 | 23.2266 | 3032.748
16045256 | 17859 | 4 | 2905E05074 | 2022-09-23T00:57:16.867+00:00 | TRUE | 4 | 0 | . | 152 | 23.2344 | 2489.225
16045256 | 17859 | 1 | 2905E04469 | 2022-09-23T01:02:07.007+00:00 | TRUE | 1 | 0 | . | 152 | 23.2422 | 3076.663
16045256 | 17859 | 4 | 2905E04661 | 2022-09-23T00:47:18.603+00:00 | FALSE | 4 | 49 | . | 152 | 23.2578 | 0
16045256 | 17859 | 1 | 2905E04786 | 2022-09-22T01:09:09.79+00:00 | TRUE | 1 | 0 | . | 152 | 23.2578 | 2925.624
16045256 | 17859 | 2 | 2905E04936 | 2022-09-22T01:39:08.163+00:00 | TRUE | 3 | 0 | . | 152 | 23.2578 | 3212.55
16045256 | 17859 | 4 | 2905E04594 | 2022-09-23T03:37:20.387+00:00 | FALSE | 4 | 71 | . | 152 | 23.2656 | 2505.483
16045256 | 17859 | 1 | 2905E04957 | 2022-09-23T00:53:55.367+00:00 | FALSE | 1 | 46 | . | 152 | 23.2891 | 0
16045256 | 17859 | 4 | 2905E04660 | 2022-09-22T01:45:54.677+00:00 | TRUE | 4 | 0 | . | 152 | 23.2969 | 2313.007
16045256 | 17859 | 3 | 2905E01253 | 2022-09-21T10:03:15.397+00:00 | FALSE | 2 | 88 | . | 152 | 23.3047 | -999
16045256 | 17859 | 1 | 2905E00972 | 2022-09-21T10:09:21.037+00:00 | FALSE | 1 | 49 | . | 152 | 23.3125 | 0
16045256 | 17859 | 1 | 2905E04949 | 2022-09-22T00:55:41.4+00:00 | TRUE | 1 | 0 | . | 152 | 23.3203 | 3010.146
16045256 | 17859 | 3 | 2905E04336 | 2022-09-22T01:12:14.737+00:00 | TRUE | 2 | 0 | . | 152 | 23.3203 | 3015.565
16045256 | 17859 | 3 | 2905E04723 | 2022-09-22T02:12:59.94+00:00 | TRUE | 2 | 0 | . | 152 | 23.3281 | 2948.661
16045256 | 17859 | 3 | 2905E04737 | 2022-09-22T00:49:40.03+00:00 | FALSE | 2 | 88 | . | 152 | 23.3438 | -999
16045256 | 17859 | 1 | 2905E04936 | 2022-09-22T00:49:51.55+00:00 | FALSE | 1 | 88 | . | 152 | 23.375 | -999
16045256 | 17859 | 4 | 2905E04957 | 2022-09-23T03:33:44.153+00:00 | TRUE | 4 | 0 | . | 152 | 23.3984 | 2541.868
16045256 | 17859 | 2 | 2905E04391 | 2022-09-22T01:49:32.657+00:00 | TRUE | 3 | 0 | . | 152 | 23.4062 | 3087.46
16045256 | 17859 | 3 | 2905E04958 | 2022-09-22T00:46:08.233+00:00 | FALSE | 2 | 88 | . | 152 | 23.4141 | -999
16045256 | 17859 | 4 | 2905E04147 | 2022-09-22T01:52:22.983+00:00 | TRUE | 4 | 0 | . | 152 | 23.4141 | 2421.349
16045256 | 17859 | 1 | 2905E01013 | 2022-09-21T10:03:52.3+00:00 | TRUE | 1 | 0 | . | 152 | 23.4375 | 2998.7
16045256 | 17859 | 4 | 2905E04689 | 2022-09-22T02:13:36.08+00:00 | TRUE | 4 | 0 | . | 152 | 23.4531 | 2426.979
16045256 | 17859 | 1 | 2905E00939 | 2022-09-21T10:12:54.93+00:00 | TRUE | 1 | 0 | . | 152 | 23.4609 | 2992.78
16045256 | 17859 | 3 | 2905E03851 | 2022-09-22T01:24:55.503+00:00 | FALSE | 2 | 88 | . | 152 | 23.4688 | -999
16045256 | 17859 | 1 | 2905E04487 | 2022-09-22T01:14:25.26+00:00 | TRUE | 1 | 0 | . | 152 | 23.4766 | 2912.048
16045256 | 17859 | 3 | 2905E01725 | 2022-09-21T10:09:34.66+00:00 | TRUE | 2 | 0 | . | 152 | 23.4844 | 2884.555
16045256 | 17859 | 3 | 2905E04327 | 2022-09-22T01:16:31.193+00:00 | FALSE | 2 | 88 | . | 152 | 23.4922 | -999
16045256 | 17859 | 3 | 2905E04480 | 2022-09-22T01:19:37.907+00:00 | TRUE | 2 | 0 | . | 152 | 23.5 | 2961.323
16045256 | 17859 | 2 | 2905E04960 | 2022-09-22T02:16:05.24+00:00 | FALSE | 3 | 88 | . | 152 | 23.5078 | -999
16045256 | 17859 | 3 | 2905E00972 | 2022-09-21T10:12:54.59+00:00 | TRUE | 2 | 0 | . | 152 | 23.5156 | 3096.254
16045256 | 17859 | 2 | 2905E04072 | 2022-09-22T02:12:21.733+00:00 | FALSE | 3 | 88 | . | 152 | 23.5312 | -999
16045256 | 17859 | 3 | 2905E03976 | 2022-09-22T01:21:24.68+00:00 | FALSE | 2 | 46 | . | 152 | 23.5391 | 0
16045256 | 17859 | 1 | 2905E04567 | 2022-09-22T01:18:36.62+00:00 | TRUE | 1 | 0 | . | 152 | 23.5391 | 2939.544
16045256 | 17859 | 3 | 2905E04019 | 2022-09-22T02:18:25.597+00:00 | TRUE | 2 | 0 | . | 152 | 23.5547 | 3067.743
16045256 | 17859 | 1 | 2905E04410 | 2022-09-22T01:11:52.523+00:00 | TRUE | 1 | 0 | . | 152 | 23.5625 | 2938.601
16045256 | 17859 | 4 | 2905E04788 | 2022-09-22T01:59:25.73+00:00 | FALSE | 4 | 76 | . | 152 | 23.5703 | 2324.601
16045256 | 17859 | 1 | 2905E04243 | 2022-09-22T00:45:55.033+00:00 | FALSE | 1 | 71 | . | 152 | 23.5703 | 3050.789
16045256 | 17859 | 2 | 2905E04477 | 2022-09-22T01:52:20.497+00:00 | TRUE | 3 | 0 | . | 152 | 23.5703 | 3154.319
16045256 | 17859 | 2 | 2905E04787 | 2022-09-22T02:20:05.323+00:00 | TRUE | 3 | 0 | . | 152 | 23.5781 | 3191.316
16045256 | 17859 | 4 | 2905E04688 | 2022-09-22T05:17:20.583+00:00 | TRUE | 4 | 0 | . | 152 | 23.5938 | 2546.991
16045256 | 17859 | 3 | 2905E01147 | 2022-09-21T10:15:53.81+00:00 | TRUE | 2 | 0 | . | 152 | 23.5938 | 3096.982
16045256 | 17859 | 2 | 2905E04564 | 2022-09-23T03:36:39.927+00:00 | TRUE | 3 | 0 | . | 152 | 23.5938 | 3154.546
16045256 | 17859 | 4 | 2905E05129 | 2022-09-22T05:25:49.01+00:00 | TRUE | 4 | 0 | . | 152 | 23.6016 | 2527.542
16045256 | 17859 | 1 | 2905E01005 | 2022-09-21T10:14:50.31+00:00 | FALSE | 1 | 49 | . | 152 | 23.6094 | 0
16045256 | 17859 | 3 | 2905E04457 | 2022-09-22T01:23:03.393+00:00 | FALSE | 2 | 88 | . | 152 | 23.625 | -999
16045256 | 17859 | 2 | 2905E04985 | 2022-09-22T02:23:00.797+00:00 | TRUE | 3 | 0 | . | 152 | 23.6328 | 3107.811
16045256 | 17859 | 2 | 2905E03852 | 2022-09-23T03:32:59.793+00:00 | TRUE | 3 | 0 | . | 152 | 23.6328 | 3164.052
16045256 | 17859 | 1 | 2905E01100 | 2022-09-21T10:17:43.173+00:00 | TRUE | 1 | 0 | . | 152 | 23.6406 | 2940.503
16045256 | 17859 | 4 | 2905E04132 | 2022-09-22T02:02:04.043+00:00 | TRUE | 4 | 0 | . | 152 | 23.6484 | 2481.864
16045256 | 17859 | 1 | 2905E03999 | 2022-09-22T01:21:13.003+00:00 | TRUE | 1 | 0 | . | 152 | 23.6484 | 2944.781
16045256 | 17859 | 4 | 2905E04594 | 2022-09-23T03:40:59.313+00:00 | FALSE | 4 | 71 | . | 152 | 23.6562 | 2501.302
16045256 | 17859 | 3 | 2905E04950 | 2022-09-22T02:15:34.26+00:00 | TRUE | 2 | 0 | . | 152 | 23.6562 | 2943.117
16045256 | 17859 | 4 | 2905E04839 | 2022-09-22T05:23:01.927+00:00 | TRUE | 4 | 0 | . | 152 | 23.6641 | 2473.145
16045256 | 17859 | 4 | 2905E03934 | 2022-09-22T01:57:46.787+00:00 | TRUE | 4 | 0 | . | 152 | 23.6641 | 2503.075
16045256 | 17859 | 4 | 2905E03811 | 2022-09-22T05:19:56.117+00:00 | TRUE | 4 | 0 | . | 152 | 23.6719 | 2475.128
16045256 | 17859 | 4 | 2905E03974 | 2022-09-22T02:41:49.397+00:00 | TRUE | 4 | 0 | . | 152 | 23.6719 | 2556.151
16045256 | 17859 | 2 | 2905E03555 | 2022-09-22T02:14:56.45+00:00 | TRUE | 3 | 0 | . | 152 | 23.6719 | 3139.982
16045256 | 17859 | 4 | 2905E04535 | 2022-09-22T05:14:50.977+00:00 | TRUE | 4 | 0 | . | 152 | 23.6875 | 2494.617
16045256 | 17859 | 1 | 2905E01259 | 2022-09-21T10:07:20.127+00:00 | TRUE | 1 | 0 | . | 152 | 23.6953 | 2996.698
16045256 | 17859 | 2 | 2905E04583 | 2022-09-22T01:57:24.34+00:00 | TRUE | 3 | 0 | . | 152 | 23.6953 | 3098.427
16045256 | 17859 | 4 | 2905E04960 | 2022-09-22T02:20:08.523+00:00 | FALSE | 4 | 88 | . | 152 | 23.7031 | -999
16045256 | 17859 | 1 | 2905E01069 | 2022-09-21T10:23:28.287+00:00 | TRUE | 1 | 0 | . | 152 | 23.7109 | 2904.833
16045256 | 17859 | 3 | 2905E04451 | 2022-09-23T03:36:34.487+00:00 | TRUE | 2 | 0 | . | 152 | 23.7266 | 2924.192
16045256 | 17859 | 3 | 2905E04035 | 2022-09-22T01:33:58.64+00:00 | FALSE | 2 | 88 | . | 152 | 23.75 | -999
16045256 | 17859 | 2 | 2905E04577 | 2022-09-22T01:54:52.033+00:00 | TRUE | 3 | 0 | . | 152 | 23.7578 | 3095.675
16045256 | 17859 | 4 | 2905E00953 | 2022-09-22T01:49:37.423+00:00 | TRUE | 4 | 0 | . | 152 | 23.7656 | 2477.851
16045256 | 17859 | 1 | 2905E01085 | 2022-09-21T10:29:23.82+00:00 | TRUE | 1 | 0 | . | 152 | 23.7812 | 3095.593
16045256 | 17859 | 1 | 2905E01291 | 2022-09-21T10:26:45.567+00:00 | TRUE | 1 | 0 | . | 152 | 23.7891 | 2942.777
16045256 | 17859 | 2 | 2905E04437 | 2022-09-22T01:59:58.753+00:00 | TRUE | 3 | 0 | . | 152 | 23.7891 | 3149.49
16045256 | 17859 | 3 | 2905E04690 | 2022-09-23T03:39:40.013+00:00 | TRUE | 2 | 0 | . | 152 | 23.7969 | 2988.543
16045256 | 17859 | 3 | 2905E04942 | 2022-09-23T03:33:32.71+00:00 | TRUE | 2 | 0 | . | 152 | 23.8047 | 2998.436
16045256 | 17859 | 3 | 2905E01253 | 2022-09-21T10:06:40.867+00:00 | TRUE | 2 | 0 | . | 152 | 23.8047 | 3064.912
16045256 | 17859 | 2 | 2905E03820 | 2022-09-22T05:18:43.603+00:00 | TRUE | 3 | 0 | . | 152 | 23.8047 | 3193.562
16045256 | 17859 | 2 | 2905E02667 | 2022-09-22T05:28:39.593+00:00 | FALSE | 3 | 88 | . | 152 | 23.8125 | -999
16045256 | 17859 | 3 | 2905E01153 | 2022-09-21T10:28:30.183+00:00 | TRUE | 2 | 0 | . | 152 | 23.8125 | 2899.564
16045256 | 17859 | 3 | 2905E03883 | 2022-09-22T01:26:44.98+00:00 | FALSE | 2 | 88 | . | 152 | 23.8203 | -999
16045256 | 17859 | 2 | 2905E04671 | 2022-09-22T02:28:00.33+00:00 | TRUE | 3 | 0 | . | 152 | 23.8203 | 3207.443
16045256 | 17859 | 3 | 2905E03942 | 2022-09-22T08:48:08.83+00:00 | FALSE | 2 | 46 | . | 152 | 23.8281 | 0
16045256 | 17859 | 4 | 2905E04036 | 2022-09-22T03:02:29.507+00:00 | TRUE | 4 | 0 | . | 152 | 23.8281 | 2497.029
16045256 | 17859 | 3 | 2905E04021 | 2022-09-22T01:29:12.547+00:00 | FALSE | 2 | 88 | . | 152 | 23.8359 | -999
16045256 | 17859 | 4 | 2905E04992 | 2022-09-22T05:29:01.883+00:00 | TRUE | 4 | 0 | . | 152 | 23.8359 | 2414.864
16045256 | 17859 | 3 | 2905E04521 | 2022-09-22T01:32:08.567+00:00 | FALSE | 2 | 88 | . | 152 | 23.8438 | -999
16045256 | 17859 | 4 | 2905E04717 | 2022-09-22T02:47:45.537+00:00 | TRUE | 4 | 0 | . | 152 | 23.8438 | 2430.812
16045256 | 17859 | 4 | 2905E05317 | 2022-09-22T02:39:13.2+00:00 | TRUE | 4 | 0 | . | 152 | 23.8438 | 2546.443
16045256 | 17859 | 4 | 2905E04584 | 2022-09-22T02:34:36.91+00:00 | TRUE | 4 | 0 | . | 152 | 23.8438 | 2576.157
16045256 | 17859 | 1 | 2905E01267 | 2022-09-22T01:32:32.853+00:00 | TRUE | 1 | 0 | . | 152 | 23.8516 | 2927.304
16045256 | 17859 | 2 | 2905E04860 | 2022-09-22T02:25:27.733+00:00 | TRUE | 3 | 0 | . | 152 | 23.8516 | 3140.405
16045256 | 17859 | 3 | 2905E05124 | 2022-09-22T05:18:40.49+00:00 | TRUE | 2 | 0 | . | 152 | 23.8594 | 3072.181
16045256 | 17859 | 2 | 2905E04139 | 2022-09-22T05:05:36.57+00:00 | FALSE | 3 | 10 | . | 152 | 23.8594 | 3164.484
16045256 | 17859 | 4 | 2905E04953 | 2022-09-22T02:36:03.427+00:00 | FALSE | 4 | 81 | . | 152 | 23.8672 | 2520.921
16045256 | 17859 | 3 | 2905E01005 | 2022-09-21T10:18:58.143+00:00 | TRUE | 2 | 0 | . | 152 | 23.8828 | 2810.458
16045256 | 17859 | 1 | 2905E04661 | 2022-09-22T05:30:15.467+00:00 | FALSE | 1 | 71 | . | 152 | 23.8828 | 3038.476
16045256 | 17859 | 1 | 2905E03951 | 2022-09-22T05:14:25.487+00:00 | TRUE | 1 | 0 | . | 152 | 23.8828 | 3085.312
16045256 | 17859 | 2 | 2905E04199 | 2022-09-22T02:39:42.673+00:00 | TRUE | 3 | 0 | . | 152 | 23.8828 | 3210.118
16045256 | 17859 | 2 | 2905E04381 | 2022-09-22T02:37:01.487+00:00 | TRUE | 3 | 0 | . | 152 | 23.8828 | 3246.831
16045256 | 17859 | 3 | 2905E04583 | 2022-09-22T01:54:25.19+00:00 | FALSE | 2 | 88 | . | 152 | 23.8906 | -999
16045256 | 17859 | 1 | 2905E01228 | 2022-09-22T01:35:12.233+00:00 | TRUE | 1 | 0 | . | 152 | 23.8906 | 2992.059
16045256 | 17859 | 1 | 2905E04349 | 2022-09-22T01:23:47.697+00:00 | TRUE | 1 | 0 | . | 152 | 23.8906 | 3019.968
16045256 | 17859 | 1 | 2905E04957 | 2022-09-22T05:18:19.363+00:00 | FALSE | 1 | 46 | . | 152 | 23.8984 | 0
16045256 | 17859 | 4 | 2905E05325 | 2022-09-22T02:44:47.78+00:00 | TRUE | 4 | 0 | . | 152 | 23.8984 | 2497.351
16045256 | 17859 | 1 | 2905E01048 | 2022-09-22T01:41:10.447+00:00 | TRUE | 1 | 0 | . | 152 | 23.8984 | 2973.938
16045256 | 17859 | 2 | 2905E04832 | 2022-09-22T05:33:25.247+00:00 | TRUE | 3 | 0 | . | 152 | 23.8984 | 3328.929
16045256 | 17859 | 3 | 2905E04258 | 2022-09-22T08:46:10.33+00:00 | FALSE | 2 | 88 | . | 152 | 23.9141 | -999
16045256 | 17859 | 3 | 2905E00996 | 2022-09-21T10:22:26.393+00:00 | TRUE | 2 | 0 | . | 152 | 23.9141 | 2882.786
16045256 | 17859 | 3 | 2905E03961 | 2022-09-22T09:46:02.437+00:00 | TRUE | 2 | 0 | . | 152 | 23.9219 | 2951.774
16045256 | 17859 | 3 | 2905E00953 | 2022-09-22T01:44:30.08+00:00 | FALSE | 2 | 104 | . | 152 | 23.9219 | 2983.799
16045256 | 17859 | 3 | 2905E03641 | 2022-09-22T09:49:57.893+00:00 | TRUE | 2 | 0 | . | 152 | 23.9219 | 3087.93
16045256 | 17859 | 2 | 2905E04956 | 2022-09-22T02:30:25.133+00:00 | TRUE | 3 | 0 | . | 152 | 23.9219 | 3178.43
16045256 | 17859 | 2 | 2905E05127 | 2022-09-22T05:15:26.88+00:00 | TRUE | 3 | 0 | . | 152 | 23.9219 | 3212.777
16045256 | 17859 | 1 | 2905E01114 | 2022-09-22T01:26:36.857+00:00 | TRUE | 1 | 0 | . | 152 | 23.9297 | 2953.351
16045256 | 17859 | 1 | 2905E04233 | 2022-09-22T05:36:40.877+00:00 | TRUE | 1 | 0 | . | 152 | 23.9297 | 3257.509
16045256 | 17859 | 3 | 2905E03705 | 2022-09-22T05:28:31.003+00:00 | FALSE | 2 | 88 | . | 152 | 23.9453 | -999
16045256 | 17859 | 1 | 2905E04850 | 2022-09-22T05:26:14.397+00:00 | TRUE | 1 | 0 | . | 152 | 23.9453 | 3001.123
16045256 | 17859 | 4 | 2905E04028 | 2022-09-22T05:31:26.863+00:00 | FALSE | 4 | 46 | . | 152 | 23.9531 | 0
16045256 | 17859 | 3 | 2905E01014 | 2022-09-21T10:31:50.787+00:00 | TRUE | 2 | 0 | . | 152 | 23.9531 | 2915.159
16045256 | 17859 | 3 | 2905E01289 | 2022-09-21T10:25:45.623+00:00 | TRUE | 2 | 0 | . | 152 | 23.9531 | 2991.583
16045256 | 17859 | 1 | 2905E01101 | 2022-09-21T10:32:09.873+00:00 | TRUE | 1 | 0 | . | 152 | 23.9531 | 3098.128
16045256 | 17859 | 1 | 2905E05114 | 2022-09-22T09:32:22.207+00:00 | TRUE | 1 | 0 | . | 152 | 23.9609 | 3059.491
16045256 | 17859 | 4 | 2905E04697 | 2022-09-22T05:36:48.113+00:00 | TRUE | 4 | 0 | . | 152 | 23.9688 | 2581.309
16045256 | 17859 | 4 | 2905E05316 | 2022-09-22T02:24:17.76+00:00 | TRUE | 4 | 0 | . | 152 | 23.9766 | 2474.631
16045256 | 17859 | 4 | 2905E04655 | 2022-09-22T02:53:11.567+00:00 | TRUE | 4 | 0 | . | 152 | 23.9844 | 2535.606
16045256 | 17859 | 3 | 2905E04958 | 2022-09-22T01:36:23.48+00:00 | FALSE | 2 | 88 | . | 152 | 24 | -999
16045256 | 17859 | 4 | 2905E04382 | 2022-09-22T02:58:31.527+00:00 | TRUE | 4 | 0 | . | 152 | 24 | 2399.716
16045256 | 17859 | 3 | 2905E04372 | 2022-09-22T02:24:00.333+00:00 | TRUE | 2 | 0 | . | 152 | 24.0078 | 3021.561
16045256 | 17859 | 1 | 2905E05336 | 2022-09-22T05:17:16.98+00:00 | TRUE | 1 | 0 | . | 152 | 24.0078 | 3027.746
16045256 | 17859 | 4 | 2905E04889 | 2022-09-22T07:25:51.317+00:00 | TRUE | 4 | 0 | . | 152 | 24.0156 | 2539.42
16045256 | 17859 | 1 | 2905E03876 | 2022-09-22T09:34:37.897+00:00 | FALSE | 1 | 88 | . | 152 | 24.0312 | -999
16045256 | 17859 | 2 | 2905E04942 | 2022-09-22T05:21:40.237+00:00 | FALSE | 3 | 88 | . | 152 | 24.0312 | -999
16045256 | 17859 | 3 | 2905E04771 | 2022-09-22T05:40:21.79+00:00 | FALSE | 2 | 88 | . | 152 | 24.0391 | -999
16045256 | 17859 | 4 | 2905E04251 | 2022-09-22T05:39:20.3+00:00 | TRUE | 4 | 0 | . | 152 | 24.0391 | 2483.015
16045256 | 17859 | 4 | 2905E04031 | 2022-09-22T05:45:12.057+00:00 | TRUE | 4 | 0 | . | 152 | 24.0391 | 2561.211
16045256 | 17859 | 4 | 2905E05110 | 2022-09-22T02:50:20.723+00:00 | TRUE | 4 | 0 | . | 152 | 24.0547 | 2490.315
16045256 | 17859 | 3 | 2905E01127 | 2022-09-21T10:37:22.867+00:00 | TRUE | 2 | 0 | . | 152 | 24.0625 | 2905.146
16045256 | 17859 | 1 | 2905E00931 | 2022-09-21T10:37:51.623+00:00 | TRUE | 1 | 0 | . | 152 | 24.0625 | 2967.096
16045256 | 17859 | 1 | 2905E00624 | 2022-09-21T10:40:26.913+00:00 | TRUE | 1 | 0 | . | 152 | 24.0625 | 2975.982
16045256 | 17859 | 3 | 2905E03661 | 2022-09-22T02:21:12.117+00:00 | TRUE | 2 | 0 | . | 152 | 24.0625 | 3013.493
16045256 | 17859 | 1 | 2905E05067 | 2022-09-21T11:09:31.54+00:00 | TRUE | 1 | 0 | . | 152 | 24.0625 | 3052.164
16045256 | 17859 | 4 | 2905E04384 | 2022-09-22T05:53:11.48+00:00 | FALSE | 4 | 88 | . | 152 | 24.0703 | -999
16045256 | 17859 | 3 | 2905E01216 | 2022-09-21T10:43:10.82+00:00 | FALSE | 2 | 46 | . | 152 | 24.0703 | 0
16045256 | 17859 | 3 | 2905E05161 | 2022-09-22T09:53:41.31+00:00 | TRUE | 2 | 0 | . | 152 | 24.0703 | 2954.523
16045256 | 17859 | 3 | 2905E01273 | 2022-09-21T10:39:56.913+00:00 | TRUE | 2 | 0 | . | 152 | 24.0703 | 2956.299
16045256 | 17859 | 4 | 2905E04825 | 2022-09-22T09:53:51.307+00:00 | TRUE | 4 | 0 | . | 152 | 24.0781 | 2569.604
16045256 | 17859 | 1 | 2905E00975 | 2022-09-21T10:34:53.913+00:00 | TRUE | 1 | 0 | . | 152 | 24.0781 | 2886.817
16045256 | 17859 | 3 | 2905E04576 | 2022-09-22T01:41:36.397+00:00 | TRUE | 2 | 0 | . | 152 | 24.0781 | 2986.078
16045256 | 17859 | 1 | 2905E05136 | 2022-09-21T11:04:13.287+00:00 | TRUE | 1 | 0 | . | 152 | 24.0781 | 3023.296
16045256 | 17859 | 2 | 2905E04897 | 2022-09-22T05:35:36.88+00:00 | FALSE | 3 | 55 | . | 152 | 24.0781 | 3189.68
16045256 | 17859 | 4 | 2905E04617 | 2022-09-22T06:05:04.897+00:00 | TRUE | 4 | 0 | . | 152 | 24.0859 | 2475.869
16045256 | 17859 | 3 | 2905E04770 | 2022-09-22T05:47:22.947+00:00 | FALSE | 2 | 88 | . | 152 | 24.0938 | -999
16045256 | 17859 | 1 | 2905E01423 | 2022-09-21T10:53:34.41+00:00 | FALSE | 1 | 88 | . | 152 | 24.0938 | -999
16045256 | 17859 | 4 | 2905E04513 | 2022-09-22T06:14:16.977+00:00 | TRUE | 4 | 0 | . | 152 | 24.0938 | 2526.496
16045256 | 17859 | 3 | 2905E04438 | 2022-09-22T10:05:34.377+00:00 | TRUE | 2 | 0 | . | 152 | 24.0938 | 2927.239
16045256 | 17859 | 3 | 2905E04714 | 2022-09-22T02:35:47.25+00:00 | TRUE | 2 | 0 | . | 152 | 24.0938 | 3079.411
16045256 | 17859 | 3 | 2905E05203 | 2022-09-22T09:58:59.897+00:00 | TRUE | 2 | 0 | . | 152 | 24.0938 | 3158.731
16045256 | 17859 | 4 | 2905E04310 | 2022-09-22T02:55:42.027+00:00 | FALSE | 4 | 88 | . | 152 | 24.1094 | -999
16045256 | 17859 | 1 | 2905E03876 | 2022-09-22T09:37:27.083+00:00 | TRUE | 1 | 0 | . | 152 | 24.1094 | 2977.519
16045256 | 17859 | 1 | 2905E04407 | 2022-09-22T01:43:42.887+00:00 | TRUE | 1 | 0 | . | 152 | 24.1094 | 2989.591
16045256 | 17859 | 1 | 2905E04243 | 2022-09-22T01:38:09.687+00:00 | TRUE | 1 | 0 | . | 152 | 24.1094 | 3019.291
16045256 | 17859 | 1 | 2905E04847 | 2022-09-21T11:01:37+00:00 | TRUE | 1 | 0 | . | 152 | 24.1094 | 3069.331
16045256 | 17859 | 4 | 2905E04780 | 2022-09-22T03:05:02.627+00:00 | TRUE | 4 | 0 | . | 152 | 24.1172 | 2491.199
16045256 | 17859 | 4 | 2905E04186 | 2022-09-22T07:23:22.097+00:00 | TRUE | 4 | 0 | . | 152 | 24.1172 | 2514.726
16045256 | 17859 | 4 | 2905E04373 | 2022-09-22T02:59:38.127+00:00 | FALSE | 4 | 55 | . | 152 | 24.1172 | 2516.566
16045256 | 17859 | 4 | 2905E04771 | 2022-09-22T06:08:36.957+00:00 | TRUE | 4 | 0 | . | 152 | 24.1172 | 2528.424
16045256 | 17859 | 3 | 2905E00948 | 2022-09-21T10:34:32.613+00:00 | TRUE | 2 | 0 | . | 152 | 24.1172 | 3010.674
16045256 | 17859 | 4 | 2905E04702 | 2022-09-22T10:07:50.653+00:00 | FALSE | 4 | 88 | . | 152 | 24.125 | -999
16045256 | 17859 | 1 | 2905E04709 | 2022-09-21T11:06:44.673+00:00 | TRUE | 1 | 0 | . | 152 | 24.125 | 2956.453
16045256 | 17859 | 1 | 2905E03772 | 2022-09-22T05:42:12.407+00:00 | TRUE | 1 | 0 | . | 152 | 24.125 | 3121.479
16045256 | 17859 | 3 | 2905E04916 | 2022-09-22T05:05:37.58+00:00 | TRUE | 2 | 0 | . | 152 | 24.1328 | 3117.117
16045256 | 17859 | 2 | 2905E03981 | 2022-09-22T05:47:44.947+00:00 | TRUE | 3 | 0 | . | 152 | 24.1328 | 3183.259
16045256 | 17859 | 3 | 2905E04577 | 2022-09-22T01:52:10.833+00:00 | FALSE | 2 | 88 | . | 152 | 24.1484 | -999
16045256 | 17859 | 1 | 2905E01216 | 2022-09-21T10:50:48.043+00:00 | FALSE | 1 | 88 | . | 152 | 24.1484 | -999
16045256 | 17859 | 4 | 2905E04048 | 2022-09-22T05:42:06.263+00:00 | TRUE | 4 | 0 | . | 152 | 24.1484 | 2448.1
16045256 | 17859 | 4 | 2905E04682 | 2022-09-22T03:25:04.063+00:00 | TRUE | 4 | 0 | . | 152 | 24.1484 | 2495.192
16045256 | 17859 | 3 | 2905E04947 | 2022-09-22T05:22:46.153+00:00 | TRUE | 2 | 0 | . | 152 | 24.1484 | 2937.749
16045256 | 17859 | 2 | 2905E04749 | 2022-09-22T05:40:27.43+00:00 | FALSE | 3 | 73 | . | 152 | 24.1484 | 3281.722
16045256 | 17859 | 4 | 2905E04641 | 2022-09-22T05:48:04.563+00:00 | TRUE | 4 | 0 | . | 152 | 24.1562 | 2436.578
16045256 | 17859 | 1 | 2905E00711 | 2022-09-21T10:56:28.003+00:00 | TRUE | 1 | 0 | . | 152 | 24.1562 | 2930.485
16045256 | 17859 | 3 | 2905E04744 | 2022-09-22T02:26:41.267+00:00 | TRUE | 2 | 0 | . | 152 | 24.1562 | 3036.83
16045256 | 17859 | 1 | 2905E05197 | 2022-09-21T11:12:03.11+00:00 | TRUE | 1 | 0 | . | 152 | 24.1641 | 2986.116
16045256 | 17859 | 1 | 2905E04436 | 2022-09-22T09:58:17.497+00:00 | TRUE | 1 | 0 | . | 152 | 24.1641 | 3059.705
16045256 | 17859 | 2 | 2905E03134 | 2022-09-22T02:34:22.08+00:00 | TRUE | 3 | 0 | . | 152 | 24.1641 | 3230.483
16045256 | 17859 | 4 | 2905E04730 | 2022-09-22T06:01:50.62+00:00 | FALSE | 4 | 88 | . | 152 | 24.1797 | -999
16045256 | 17859 | 4 | 2905E04330 | 2022-09-22T05:51:01.293+00:00 | TRUE | 4 | 0 | . | 152 | 24.1797 | 2568.969
16045256 | 17859 | 3 | 2905E05285 | 2022-09-22T05:15:09.547+00:00 | TRUE | 2 | 0 | . | 152 | 24.1797 | 2991.386
16045256 | 17859 | 1 | 2905E04622 | 2022-09-22T05:39:23.29+00:00 | TRUE | 1 | 0 | . | 152 | 24.1797 | 3036.791
16045256 | 17859 | 2 | 2905E04831 | 2022-09-22T02:56:26.893+00:00 | TRUE | 3 | 0 | . | 152 | 24.1797 | 3175.059
16045256 | 17859 | 1 | 2905E04983 | 2022-09-22T10:05:11.507+00:00 | TRUE | 1 | 0 | . | 152 | 24.1875 | 2996.353
16045256 | 17859 | 1 | 2905E04842 | 2022-09-22T07:44:57.28+00:00 | FALSE | 1 | 88 | . | 152 | 24.1953 | -999
16045256 | 17859 | 4 | 2905E04521 | 2022-09-22T04:27:28.667+00:00 | FALSE | 4 | 79 | . | 152 | 24.1953 | 2418.947
16045256 | 17859 | 4 | 2905E05185 | 2022-09-22T07:19:55.54+00:00 | TRUE | 4 | 0 | . | 152 | 24.1953 | 2454.135
16045256 | 17859 | 4 | 2905E05228 | 2022-09-22T07:28:06+00:00 | FALSE | 4 | 76 | . | 152 | 24.1953 | 2555.075
16045256 | 17859 | 4 | 2905E04754 | 2022-09-22T08:22:59.17+00:00 | TRUE | 4 | 0 | . | 152 | 24.1953 | 2578.063
16045256 | 17859 | 4 | 2905E04589 | 2022-09-22T03:20:42.433+00:00 | TRUE | 4 | 0 | . | 152 | 24.1953 | 2578.583
16045256 | 17859 | 3 | 2905E04943 | 2022-09-22T10:01:32.08+00:00 | TRUE | 2 | 0 | . | 152 | 24.1953 | 3007.985
16045256 | 17859 | 1 | 2905E04271 | 2022-09-22T10:00:47.01+00:00 | TRUE | 1 | 0 | . | 152 | 24.1953 | 3144.711
16045256 | 17859 | 4 | 2905E05259 | 2022-09-22T07:39:08.863+00:00 | FALSE | 4 | 88 | . | 152 | 24.2031 | -999
16045256 | 17859 | 3 | 2905E04720 | 2022-09-22T10:07:54.7+00:00 | TRUE | 2 | 0 | . | 152 | 24.2031 | 3049.343
16045256 | 17859 | 2 | 2905E04897 | 2022-09-22T06:08:18.533+00:00 | TRUE | 3 | 0 | . | 152 | 24.2031 | 3219.983
16045256 | 17859 | 3 | 2905E04169 | 2022-09-22T05:35:39.603+00:00 | FALSE | 2 | 88 | . | 152 | 24.2109 | -999
16045256 | 17859 | 4 | 2905E04872 | 2022-09-22T10:49:13.527+00:00 | TRUE | 4 | 0 | . | 152 | 24.2109 | 2477.837
16045256 | 17859 | 4 | 2905E03586 | 2022-09-22T03:38:31.487+00:00 | TRUE | 4 | 0 | . | 152 | 24.2109 | 2564.902
16045256 | 17859 | 1 | 2905E05243 | 2022-09-22T02:19:32.67+00:00 | TRUE | 1 | 0 | . | 152 | 24.2109 | 2945.306
16045256 | 17859 | 3 | 2905E03981 | 2022-09-22T05:42:55.58+00:00 | FALSE | 2 | 88 | . | 152 | 24.2188 | -999
16045256 | 17859 | 4 | 2905E03691 | 2022-09-22T04:14:56.57+00:00 | TRUE | 4 | 0 | . | 152 | 24.2188 | 2444.266
16045256 | 17859 | 4 | 2905E03693 | 2022-09-22T04:08:46.707+00:00 | TRUE | 4 | 0 | . | 152 | 24.2266 | 2443.152
16045256 | 17859 | 4 | 2905E03991 | 2022-09-22T10:55:21.883+00:00 | TRUE | 4 | 0 | . | 152 | 24.2266 | 2519.591
16045256 | 17859 | 4 | 2905E05029 | 2022-09-22T10:43:35.567+00:00 | TRUE | 4 | 0 | . | 152 | 24.2266 | 2580.924
16045256 | 17859 | 1 | 2905E01142 | 2022-09-22T01:46:17.12+00:00 | TRUE | 1 | 0 | . | 152 | 24.2266 | 2976.581
16045256 | 17859 | 3 | 2905E04134 | 2022-09-22T02:28:31.993+00:00 | FALSE | 2 | 77 | . | 152 | 24.2266 | 2994.068
16045256 | 17859 | 4 | 2905E04531 | 2022-09-22T05:56:42.117+00:00 | FALSE | 4 | 88 | . | 152 | 24.2344 | -999
16045256 | 17859 | 3 | 2905E05074 | 2022-09-22T05:24:46.51+00:00 | FALSE | 2 | 88 | . | 152 | 24.2344 | -999
16045256 | 17859 | 4 | 2905E04765 | 2022-09-22T10:22:16.883+00:00 | TRUE | 4 | 0 | . | 152 | 24.2344 | 2591.671
16045256 | 17859 | 3 | 2905E03739 | 2022-09-22T02:55:15.753+00:00 | TRUE | 2 | 0 | . | 152 | 24.2344 | 2943.16
16045256 | 17859 | 4 | 2905E04811 | 2022-09-22T10:13:12.82+00:00 | FALSE | 4 | 88 | . | 152 | 24.2422 | -999
16045256 | 17859 | 1 | 2905E04436 | 2022-09-22T09:50:17.367+00:00 | FALSE | 1 | 49 | . | 152 | 24.2422 | 0
16045256 | 17859 | 4 | 2905E04471 | 2022-09-22T11:15:56.717+00:00 | TRUE | 4 | 0 | . | 152 | 24.2422 | 2553.847
16045256 | 17859 | 4 | 2905E04726 | 2022-09-22T07:44:45.15+00:00 | TRUE | 4 | 0 | . | 152 | 24.2422 | 2600.51
16045256 | 17859 | 4 | 2905E04375 | 2022-09-22T10:19:56.407+00:00 | TRUE | 4 | 0 | . | 152 | 24.2422 | 2612.557
16045256 | 17859 | 4 | 2905E04706 | 2022-09-22T06:11:38.083+00:00 | TRUE | 4 | 0 | . | 152 | 24.25 | 2419.359
16045256 | 17859 | 4 | 2905E04635 | 2022-09-22T03:18:17.293+00:00 | TRUE | 4 | 0 | . | 152 | 24.25 | 2577.983
16045256 | 17859 | 3 | 2905E00925 | 2022-09-21T10:52:07.907+00:00 | TRUE | 2 | 0 | . | 152 | 24.25 | 2994.757
16045256 | 17859 | 2 | 2905E04013 | 2022-09-22T09:59:19.157+00:00 | TRUE | 3 | 0 | . | 152 | 24.25 | 3158.653
16045256 | 17859 | 2 | 2905E05030 | 2022-09-22T09:50:14.877+00:00 | TRUE | 3 | 0 | . | 152 | 24.25 | 3258.175
16045256 | 17859 | 4 | 2905E05096 | 2022-09-22T07:42:19.023+00:00 | TRUE | 4 | 0 | . | 152 | 24.2578 | 2568.835
16045256 | 17859 | 4 | 2905E04072 | 2022-09-22T02:16:19.713+00:00 | TRUE | 4 | 0 | . | 152 | 24.2656 | 2411.827
16045256 | 17859 | 4 | 2905E04077 | 2022-09-22T10:12:10.05+00:00 | TRUE | 4 | 0 | . | 152 | 24.2656 | 2569.567
16045256 | 17859 | 1 | 2905E03972 | 2022-09-22T05:47:25.983+00:00 | TRUE | 1 | 0 | . | 152 | 24.2656 | 3011.505
16045256 | 17859 | 4 | 2905E04540 | 2022-09-22T06:32:20.377+00:00 | FALSE | 4 | 75 | . | 152 | 24.2734 | 2455.963
16045256 | 17859 | 4 | 2905E03862 | 2022-09-22T08:11:18.41+00:00 | TRUE | 4 | 0 | . | 152 | 24.2734 | 2472.321
16045256 | 17859 | 4 | 2905E04528 | 2022-09-22T07:32:34.483+00:00 | FALSE | 4 | 78 | . | 152 | 24.2734 | 2640.281
16045256 | 17859 | 3 | 2905E01354 | 2022-09-21T10:55:28.61+00:00 | TRUE | 2 | 0 | . | 152 | 24.2734 | 2811.986
16045256 | 17859 | 1 | 2905E04224 | 2022-09-22T01:51:23.017+00:00 | TRUE | 1 | 0 | . | 152 | 24.2734 | 2966.369
16045256 | 17859 | 2 | 2905E03700 | 2022-09-22T09:56:32.543+00:00 | TRUE | 3 | 0 | . | 152 | 24.2734 | 2976
16045256 | 17859 | 3 | 2905E05108 | 2022-09-22T02:40:06.347+00:00 | TRUE | 2 | 0 | . | 152 | 24.2734 | 3099.239
16045256 | 17859 | 3 | 2905E03885 | 2022-09-22T09:56:15.24+00:00 | TRUE | 2 | 0 | . | 152 | 24.2734 | 3126.875
16045256 | 17859 | 2 | 2905E03913 | 2022-09-22T02:52:56.16+00:00 | TRUE | 3 | 0 | . | 152 | 24.2734 | 3226.947
16045256 | 17859 | 4 | 2905E04811 | 2022-09-22T10:39:40.967+00:00 | FALSE | 4 | 88 | . | 152 | 24.2812 | -999
16045256 | 17859 | 3 | 2905E00861 | 2022-09-21T10:41:32.867+00:00 | FALSE | 2 | 88 | . | 152 | 24.2812 | -999
16045256 | 17859 | 4 | 2905E04677 | 2022-09-22T07:37:30.343+00:00 | FALSE | 4 | 112 | . | 152 | 24.2812 | 2412.543
16045256 | 17859 | 4 | 2905E03689 | 2022-09-22T06:30:38.333+00:00 | TRUE | 4 | 0 | . | 152 | 24.2812 | 2502.469
16045256 | 17859 | 4 | 2905E04873 | 2022-09-22T10:58:43.213+00:00 | TRUE | 4 | 0 | . | 152 | 24.2812 | 2549.309
16045256 | 17859 | 1 | 2905E03931 | 2022-09-22T01:48:58.6+00:00 | TRUE | 1 | 0 | . | 152 | 24.2812 | 2899.412
16045256 | 17859 | 3 | 2905E01423 | 2022-09-21T10:58:14.937+00:00 | TRUE | 2 | 0 | . | 152 | 24.2812 | 2984.323
16045256 | 17859 | 1 | 2905E04054 | 2022-09-22T02:16:47.71+00:00 | TRUE | 1 | 0 | . | 152 | 24.2812 | 3031.728
16045256 | 17859 | 2 | 2905E04001 | 2022-09-22T05:44:08.337+00:00 | TRUE | 3 | 0 | . | 152 | 24.2812 | 3244.671
16045256 | 17859 | 3 | 2905E03924 | 2022-09-22T05:50:40.773+00:00 | FALSE | 2 | 88 | . | 152 | 24.2891 | -999
16045256 | 17859 | 4 | 2905E04704 | 2022-09-22T10:02:36.377+00:00 | TRUE | 4 | 0 | . | 152 | 24.2891 | 2521.677
16045256 | 17859 | 4 | 2905E03924 | 2022-09-22T06:27:47.23+00:00 | TRUE | 4 | 0 | . | 152 | 24.2891 | 2542.514
16045256 | 17859 | 3 | 2905E05292 | 2022-09-21T11:02:05.967+00:00 | TRUE | 2 | 0 | . | 152 | 24.2891 | 2942.646
16045256 | 17859 | 3 | 2905E03895 | 2022-09-22T02:42:10.52+00:00 | FALSE | 2 | 59 | . | 152 | 24.2891 | 3060.809
16045256 | 17859 | 2 | 2905E03895 | 2022-09-22T02:49:30.623+00:00 | TRUE | 3 | 0 | . | 152 | 24.2891 | 3188.78
16045256 | 17859 | 2 | 2905E04982 | 2022-09-22T02:46:08.593+00:00 | FALSE | 3 | 34 | . | 152 | 24.2891 | 3249.287
16045256 | 17859 | 4 | 2905E04504 | 2022-09-22T03:12:55.027+00:00 | TRUE | 4 | 0 | . | 152 | 24.2969 | 2459.459
16045256 | 17859 | 1 | 2905E04486 | 2022-09-22T05:44:59.407+00:00 | TRUE | 1 | 0 | . | 152 | 24.2969 | 2975.625
16045256 | 17859 | 3 | 2905E04678 | 2022-09-21T11:10:37.73+00:00 | TRUE | 2 | 0 | . | 152 | 24.2969 | 3027.635
16045256 | 17859 | 3 | 2905E04069 | 2022-09-22T10:13:41.417+00:00 | TRUE | 2 | 0 | . | 152 | 24.2969 | 3112.24
16045256 | 17859 | 1 | 2905E04371 | 2022-09-22T08:49:30.307+00:00 | TRUE | 1 | 0 | . | 152 | 24.2969 | 3172.74
16045256 | 17859 | 2 | 2905E04624 | 2022-09-22T06:16:30.387+00:00 | TRUE | 3 | 0 | . | 152 | 24.2969 | 3215.055
16045256 | 17859 | 4 | 2905E03951 | 2022-09-22T03:26:40.633+00:00 | FALSE | 4 | 88 | . | 152 | 24.3047 | -999
16045256 | 17859 | 4 | 2905E03969 | 2022-09-22T03:16:01.287+00:00 | TRUE | 4 | 0 | . | 152 | 24.3047 | 2488.4
16045256 | 17859 | 4 | 2905E04142 | 2022-09-22T07:56:17.373+00:00 | TRUE | 4 | 0 | . | 152 | 24.3047 | 2561.286
16045256 | 17859 | 4 | 2905E04010 | 2022-09-22T10:17:20.62+00:00 | TRUE | 4 | 0 | . | 152 | 24.3125 | 2626.496
16045256 | 17859 | 2 | 2905E04594 | 2022-09-22T06:33:23.367+00:00 | FALSE | 3 | 71 | . | 152 | 24.3125 | 3232.979
16045256 | 17859 | 4 | 2905E03980 | 2022-09-22T07:13:51.197+00:00 | FALSE | 4 | 88 | . | 152 | 24.3203 | -999
16045256 | 17859 | 2 | 2905E04298 | 2022-09-22T06:13:29.75+00:00 | FALSE | 3 | 88 | . | 152 | 24.3203 | -999
16045256 | 17859 | 1 | 2905E00861 | 2022-09-21T10:44:40.323+00:00 | FALSE | 1 | 59 | . | 152 | 24.3203 | 2772.144
16045256 | 17859 | 1 | 2905E04822 | 2022-09-22T02:22:42.377+00:00 | TRUE | 1 | 0 | . | 152 | 24.3203 | 2948.994
16045256 | 17859 | 1 | 2905E04578 | 2022-09-22T09:54:41.403+00:00 | TRUE | 1 | 0 | . | 152 | 24.3203 | 3028.478
16045256 | 17859 | 4 | 2905E04290 | 2022-09-22T09:59:33.07+00:00 | FALSE | 4 | 88 | . | 152 | 24.3281 | -999
16045256 | 17859 | 1 | 2905E04469 | 2022-09-22T07:22:30.183+00:00 | FALSE | 1 | 88 | . | 152 | 24.3281 | -999
16045256 | 17859 | 4 | 2905E04101 | 2022-09-22T06:15:51.85+00:00 | FALSE | 4 | 88 | . | 152 | 24.3281 | -999
16045256 | 17859 | 4 | 2905E05189 | 2022-09-22T03:33:08.403+00:00 | TRUE | 4 | 0 | . | 152 | 24.3281 | 2460.912
16045256 | 17859 | 4 | 2905E04916 | 2022-09-22T03:29:44.457+00:00 | FALSE | 4 | 10 | . | 152 | 24.3281 | 2593.962
16045256 | 17859 | 1 | 2905E04898 | 2022-09-22T05:50:24.123+00:00 | TRUE | 1 | 0 | . | 152 | 24.3281 | 2998.417
16045256 | 17859 | 2 | 2905E03888 | 2022-09-22T07:37:43.11+00:00 | TRUE | 3 | 0 | . | 152 | 24.3281 | 3144.263
16045256 | 17859 | 1 | 2905E04653 | 2022-09-22T07:26:20.273+00:00 | TRUE | 1 | 0 | . | 152 | 24.3281 | 3155.282
16045256 | 17859 | 2 | 2905E05048 | 2022-09-22T09:47:34.31+00:00 | TRUE | 3 | 0 | . | 152 | 24.3281 | 3158.763
16045256 | 17859 | 4 | 2905E05120 | 2022-09-22T10:04:10.967+00:00 | FALSE | 4 | 88 | . | 152 | 24.3359 | -999
16045256 | 17859 | 1 | 2905E04751 | 2022-09-22T07:19:51.953+00:00 | FALSE | 1 | 88 | . | 152 | 24.3359 | -999
16045256 | 17859 | 4 | 2905E03714 | 2022-09-22T03:35:37.773+00:00 | TRUE | 4 | 0 | . | 152 | 24.3359 | 2477.441
16045256 | 17859 | 2 | 2905E04676 | 2022-09-22T09:53:36.24+00:00 | TRUE | 3 | 0 | . | 152 | 24.3359 | 3343.661
16045256 | 17859 | 1 | 2905E04984 | 2022-09-22T10:07:35.563+00:00 | TRUE | 1 | 0 | . | 152 | 24.3438 | 3084.124
16045256 | 17859 | 2 | 2905E04531 | 2022-09-22T05:59:26.177+00:00 | FALSE | 3 | 88 | . | 152 | 24.3516 | -999
16045256 | 17859 | 4 | 2905E04520 | 2022-09-22T04:25:17.053+00:00 | TRUE | 4 | 0 | . | 152 | 24.3516 | 2474.037
16045256 | 17859 | 4 | 2905E04095 | 2022-09-22T11:13:20.793+00:00 | TRUE | 4 | 0 | . | 152 | 24.3516 | 2513.332
16045256 | 17859 | 1 | 2905E05228 | 2022-09-22T07:32:44.71+00:00 | TRUE | 1 | 0 | . | 152 | 24.3516 | 3105.895
16045256 | 17859 | 1 | 2905E04528 | 2022-09-22T07:35:22.1+00:00 | TRUE | 1 | 0 | . | 152 | 24.3516 | 3198.276
16045256 | 17859 | 4 | 2905E00861 | 2022-09-22T11:01:04.48+00:00 | TRUE | 4 | 0 | . | 152 | 24.3594 | 2297.389
16045256 | 17859 | 3 | 2905E04805 | 2022-09-21T11:04:59.113+00:00 | TRUE | 2 | 0 | . | 152 | 24.3594 | 2925.411
16045256 | 17859 | 1 | 2905E01363 | 2022-09-21T10:43:14.053+00:00 | TRUE | 1 | 0 | . | 152 | 24.3594 | 2976.698
16045256 | 17859 | 1 | 2905E03943 | 2022-09-22T06:13:45.473+00:00 | TRUE | 1 | 0 | . | 152 | 24.3594 | 3025.332
16045256 | 17859 | 1 | 2905E04840 | 2022-09-22T07:59:08.783+00:00 | TRUE | 1 | 0 | . | 152 | 24.3594 | 3188.278
16045256 | 17859 | 2 | 2905E03868 | 2022-09-22T05:55:33.637+00:00 | TRUE | 3 | 0 | . | 152 | 24.3594 | 3203.149
16045256 | 17859 | 2 | 2905E03922 | 2022-09-22T10:01:44.727+00:00 | TRUE | 3 | 0 | . | 152 | 24.3594 | 3326.232
16045256 | 17859 | 4 | 2905E04866 | 2022-09-22T09:55:25.547+00:00 | FALSE | 4 | 46 | . | 152 | 24.3672 | 0
16045256 | 17859 | 1 | 2905E04866 | 2022-09-22T09:44:16.533+00:00 | FALSE | 1 | 46 | . | 152 | 24.3672 | 0
16045256 | 17859 | 4 | 2905E03649 | 2022-09-22T11:04:14.91+00:00 | TRUE | 4 | 0 | . | 152 | 24.3672 | 2577.585
16045256 | 17859 | 3 | 2905E01423 | 2022-09-21T10:46:26.877+00:00 | TRUE | 2 | 0 | . | 152 | 24.3672 | 2991.924
16045256 | 17859 | 4 | 2905E03883 | 2022-09-22T04:30:27.327+00:00 | TRUE | 4 | 0 | . | 152 | 24.375 | 2500.238
16045256 | 17859 | 1 | 2905E05022 | 2022-09-22T02:27:36.723+00:00 | TRUE | 1 | 0 | . | 152 | 24.375 | 2995.142
16045256 | 17859 | 1 | 2905E04662 | 2022-09-22T10:18:53.227+00:00 | FALSE | 1 | 88 | . | 152 | 24.3828 | -999
16045256 | 17859 | 2 | 2905E04384 | 2022-09-22T05:57:33.977+00:00 | FALSE | 3 | 88 | . | 152 | 24.3828 | -999
16045256 | 17859 | 3 | 2905E04647 | 2022-09-21T11:13:22.117+00:00 | TRUE | 2 | 0 | . | 152 | 24.3828 | 3048.577
16045256 | 17859 | 2 | 2905E04524 | 2022-09-22T02:59:10.72+00:00 | TRUE | 3 | 0 | . | 152 | 24.3828 | 3201.451
16045256 | 17859 | 4 | 2905E03968 | 2022-09-22T08:34:13.667+00:00 | TRUE | 4 | 0 | . | 152 | 24.3906 | 2500.398
16045256 | 17859 | 1 | 2905E04265 | 2022-09-22T02:14:23.6+00:00 | TRUE | 1 | 0 | . | 152 | 24.3906 | 2994.061
16045256 | 17859 | 2 | 2905E04710 | 2022-09-22T03:06:32.63+00:00 | TRUE | 3 | 0 | . | 152 | 24.3906 | 3148.461
16045256 | 17859 | 4 | 2905E05097 | 2022-09-22T04:20:20.13+00:00 | TRUE | 4 | 0 | . | 152 | 24.3984 | 2434.117
16045256 | 17859 | 4 | 2905E04221 | 2022-09-22T03:56:04.417+00:00 | TRUE | 4 | 0 | . | 152 | 24.3984 | 2562.379
16045256 | 17859 | 3 | 2905E04559 | 2022-09-22T02:50:00.393+00:00 | TRUE | 2 | 0 | . | 152 | 24.3984 | 3019.312
16045256 | 17859 | 1 | 2905E04042 | 2022-09-22T06:00:50.6+00:00 | TRUE | 1 | 0 | . | 152 | 24.3984 | 3055.876
16045256 | 17859 | 1 | 2905E04365 | 2022-09-22T07:51:49.363+00:00 | TRUE | 1 | 0 | . | 152 | 24.3984 | 3058.157
16045256 | 17859 | 1 | 2905E04207 | 2022-09-22T02:43:03.267+00:00 | TRUE | 1 | 0 | . | 152 | 24.3984 | 3078.535
16045256 | 17859 | 4 | 2905E04695 | 2022-09-22T07:48:08.757+00:00 | TRUE | 4 | 0 | . | 152 | 24.4062 | 2500.186
16045256 | 17859 | 4 | 2905E04137 | 2022-09-22T06:18:21.677+00:00 | TRUE | 4 | 0 | . | 152 | 24.4062 | 2505.301
16045256 | 17859 | 4 | 2905E05296 | 2022-09-22T10:37:06.597+00:00 | TRUE | 4 | 0 | . | 152 | 24.4062 | 2621.115
16045256 | 17859 | 2 | 2905E05249 | 2022-09-22T10:09:08.863+00:00 | TRUE | 3 | 0 | . | 152 | 24.4062 | 3131.782
16045256 | 17859 | 3 | 2905E04585 | 2022-09-22T10:17:35.8+00:00 | TRUE | 2 | 0 | . | 152 | 24.4141 | 2944.126
16045256 | 17859 | 1 | 2905E03942 | 2022-09-22T09:40:55.357+00:00 | TRUE | 1 | 0 | . | 152 | 24.4141 | 3095.012
16045256 | 17859 | 2 | 2905E04942 | 2022-09-22T06:17:53.333+00:00 | FALSE | 3 | 88 | . | 152 | 24.4219 | -999
16045256 | 17859 | 4 | 2905E04879 | 2022-09-22T10:29:51.123+00:00 | TRUE | 4 | 0 | . | 152 | 24.4219 | 2480.363
16045256 | 17859 | 4 | 2905E04777 | 2022-09-22T08:20:14.327+00:00 | TRUE | 4 | 0 | . | 152 | 24.4219 | 2518.752
16045256 | 17859 | 1 | 2905E03882 | 2022-09-22T01:59:31.05+00:00 | TRUE | 1 | 0 | . | 152 | 24.4219 | 2836.448
16045256 | 17859 | 1 | 2905E03875 | 2022-09-22T10:12:43.623+00:00 | TRUE | 1 | 0 | . | 152 | 24.4219 | 3011.275
16045256 | 17859 | 1 | 2905E04835 | 2022-09-22T10:21:36.727+00:00 | TRUE | 1 | 0 | . | 152 | 24.4219 | 3114.103
16045256 | 17859 | 1 | 2905E04608 | 2022-09-22T10:24:41.517+00:00 | TRUE | 1 | 0 | . | 152 | 24.4219 | 3132.664
16045256 | 17859 | 2 | 2905E04526 | 2022-09-22T07:28:43.66+00:00 | TRUE | 3 | 0 | . | 152 | 24.4219 | 3294.485
16045256 | 17859 | 3 | 2905E04828 | 2022-09-22T02:52:50.903+00:00 | TRUE | 2 | 0 | . | 152 | 24.4297 | 3097.768
16045256 | 17859 | 2 | 2905E05284 | 2022-09-22T07:23:08.873+00:00 | TRUE | 3 | 0 | . | 152 | 24.4297 | 3350.06
16045256 | 17859 | 1 | 2905E03952 | 2022-09-22T10:02:12.647+00:00 | FALSE | 1 | 88 | . | 152 | 24.4375 | -999
16045256 | 17859 | 4 | 2905E05074 | 2022-09-22T06:20:16.057+00:00 | FALSE | 4 | 88 | . | 152 | 24.4375 | -999
16045256 | 17859 | 1 | 2905E04866 | 2022-09-22T09:47:06.913+00:00 | FALSE | 1 | 46 | . | 152 | 24.4375 | 0
16045256 | 17859 | 4 | 2905E03695 | 2022-09-22T10:06:45.42+00:00 | TRUE | 4 | 0 | . | 152 | 24.4375 | 2408.409
16045256 | 17859 | 4 | 2905E04719 | 2022-09-22T06:55:01.767+00:00 | TRUE | 4 | 0 | . | 152 | 24.4375 | 2461.65
16045256 | 17859 | 4 | 2905E05084 | 2022-09-22T04:17:35.307+00:00 | TRUE | 4 | 0 | . | 152 | 24.4375 | 2480.079
16045256 | 17859 | 1 | 2905E04745 | 2022-09-22T05:53:49.03+00:00 | TRUE | 1 | 0 | . | 152 | 24.4375 | 3040.53
16045256 | 17859 | 2 | 2905E04770 | 2022-09-22T05:50:50.927+00:00 | TRUE | 3 | 0 | . | 152 | 24.4375 | 3217.426
16045256 | 17859 | 4 | 2905E03938 | 2022-09-22T08:03:08.823+00:00 | FALSE | 4 | 88 | . | 152 | 24.4453 | -999
16045256 | 17859 | 1 | 2905E04872 | 2022-09-22T07:14:06.077+00:00 | FALSE | 1 | 88 | . | 152 | 24.4453 | -999
16045256 | 17859 | 4 | 2905E04637 | 2022-09-22T03:09:36.51+00:00 | TRUE | 4 | 0 | . | 152 | 24.4453 | 2527.893
16045256 | 17859 | 1 | 2905E05028 | 2022-09-22T02:25:14.427+00:00 | TRUE | 1 | 0 | . | 152 | 24.4453 | 2938.728
16045256 | 17859 | 2 | 2905E04571 | 2022-09-22T07:46:00.583+00:00 | TRUE | 3 | 0 | . | 152 | 24.4453 | 3328.034
16045256 | 17859 | 4 | 2905E04632 | 2022-09-22T10:33:36.843+00:00 | FALSE | 4 | 88 | . | 152 | 24.4531 | -999
16045256 | 17859 | 1 | 2905E04632 | 2022-09-22T10:15:28.767+00:00 | FALSE | 1 | 88 | . | 152 | 24.4531 | -999
16045256 | 17859 | 1 | 2905E04736 | 2022-09-22T07:28:10.827+00:00 | FALSE | 1 | 88 | . | 152 | 24.4531 | -999
16045256 | 17859 | 3 | 2905E05029 | 2022-09-22T10:10:49.027+00:00 | FALSE | 2 | 46 | . | 152 | 24.4531 | 0
16045256 | 17859 | 4 | 2905E04605 | 2022-09-22T06:36:37.42+00:00 | TRUE | 4 | 0 | . | 152 | 24.4531 | 2486.12
16045256 | 17859 | 4 | 2905E04952 | 2022-09-22T08:16:32.963+00:00 | FALSE | 4 | 74 | . | 152 | 24.4531 | 2514.199
16045256 | 17859 | 4 | 2905E04465 | 2022-09-22T08:08:57.34+00:00 | TRUE | 4 | 0 | . | 152 | 24.4531 | 2537.164
16045256 | 17859 | 2 | 2905E05194 | 2022-09-22T10:25:03.01+00:00 | TRUE | 3 | 0 | . | 152 | 24.4531 | 3235.191
16045256 | 17859 | 4 | 2905E04141 | 2022-09-22T09:58:04.413+00:00 | TRUE | 4 | 0 | . | 152 | 24.4609 | 2480.982
16045256 | 17859 | 4 | 2905E04100 | 2022-09-22T07:53:36.713+00:00 | TRUE | 4 | 0 | . | 152 | 24.4609 | 2534.516
16045256 | 17859 | 1 | 2905E04788 | 2022-09-22T02:02:08.443+00:00 | TRUE | 1 | 0 | . | 152 | 24.4609 | 2905.454
16045256 | 17859 | 1 | 2905E04002 | 2022-09-22T01:57:00.107+00:00 | TRUE | 1 | 0 | . | 152 | 24.4609 | 2989.74
16045256 | 17859 | 2 | 2905E04254 | 2022-09-22T07:40:24.797+00:00 | TRUE | 3 | 0 | . | 152 | 24.4609 | 3239.762
16045256 | 17859 | 2 | 2905E04028 | 2022-09-22T06:27:03.707+00:00 | TRUE | 3 | 0 | . | 152 | 24.4609 | 3240.977
16045256 | 17859 | 1 | 2905E05105 | 2022-09-22T10:17:22.823+00:00 | FALSE | 1 | 88 | . | 152 | 24.4688 | -999
16045256 | 17859 | 1 | 2905E03980 | 2022-09-22T07:16:20.35+00:00 | FALSE | 1 | 88 | . | 152 | 24.4688 | -999
16045256 | 17859 | 3 | 2905E05138 | 2022-09-22T02:47:14.647+00:00 | TRUE | 2 | 0 | . | 152 | 24.4688 | 3026.785
16045256 | 17859 | 2 | 2905E04846 | 2022-09-22T03:16:51.113+00:00 | TRUE | 3 | 0 | . | 152 | 24.4688 | 3214.021
16045256 | 17859 | 1 | 2905E04489 | 2022-09-22T08:08:59.2+00:00 | FALSE | 1 | 88 | . | 152 | 24.4766 | -999
16045256 | 17859 | 4 | 2905E04489 | 2022-09-22T08:15:03.89+00:00 | TRUE | 4 | 0 | . | 152 | 24.4766 | 2382.423
16045256 | 17859 | 1 | 2905E04631 | 2022-09-22T02:39:04.307+00:00 | TRUE | 1 | 0 | . | 152 | 24.4766 | 3094.19
16045256 | 17859 | 2 | 2905E04531 | 2022-09-22T06:21:43.177+00:00 | TRUE | 3 | 0 | . | 152 | 24.4766 | 3193.962
16045256 | 17859 | 4 | 2905E03851 | 2022-09-22T06:46:20.07+00:00 | TRUE | 4 | 0 | . | 152 | 24.4844 | 2437.405
16045256 | 17859 | 4 | 2905E04144 | 2022-09-22T06:52:12.407+00:00 | TRUE | 4 | 0 | . | 152 | 24.4844 | 2466.798
16045256 | 17859 | 4 | 2905E04139 | 2022-09-22T03:41:32.623+00:00 | FALSE | 4 | 10 | . | 152 | 24.4844 | 2569.68
16045256 | 17859 | 2 | 2905E04027 | 2022-09-22T03:01:38.4+00:00 | TRUE | 3 | 0 | . | 152 | 24.4844 | 2966.588
16045256 | 17859 | 1 | 2905E04274 | 2022-09-22T08:06:54.697+00:00 | TRUE | 1 | 0 | . | 152 | 24.4844 | 3012.89
16045256 | 17859 | 1 | 2905E05320 | 2022-09-22T02:52:13.193+00:00 | TRUE | 1 | 0 | . | 152 | 24.4844 | 3047.282
16045256 | 17859 | 1 | 2905E04169 | 2022-09-22T06:07:56.787+00:00 | TRUE | 1 | 0 | . | 152 | 24.4844 | 3076.249
16045256 | 17859 | 1 | 2905E03133 | 2022-09-22T09:02:27.193+00:00 | TRUE | 1 | 0 | . | 152 | 24.4844 | 3080.151
16045256 | 17859 | 1 | 2905E04101 | 2022-09-22T06:18:57.683+00:00 | TRUE | 1 | 0 | . | 152 | 24.4844 | 3091.857
16045256 | 17859 | 2 | 2905E04731 | 2022-09-22T03:27:43.543+00:00 | TRUE | 3 | 0 | . | 152 | 24.4844 | 3151.757
16045256 | 17859 | 1 | 2905E04564 | 2022-09-22T08:00:27.897+00:00 | FALSE | 1 | 88 | . | 152 | 24.4922 | -999
16045256 | 17859 | 4 | 2905E04773 | 2022-09-22T11:10:23.633+00:00 | TRUE | 4 | 0 | . | 152 | 24.4922 | 2504.532
16045256 | 17859 | 1 | 2905E04834 | 2022-09-22T06:34:30.363+00:00 | TRUE | 1 | 0 | . | 152 | 24.4922 | 3022.42
16045256 | 17859 | 1 | 2905E04675 | 2022-09-22T10:27:09.153+00:00 | TRUE | 1 | 0 | . | 152 | 24.4922 | 3113.192
16045256 | 17859 | 1 | 2905E04842 | 2022-09-22T07:46:20.617+00:00 | FALSE | 1 | 88 | . | 152 | 24.5 | -999
16045256 | 17859 | 4 | 2905E04989 | 2022-09-22T03:48:08.797+00:00 | TRUE | 4 | 0 | . | 152 | 24.5 | 2527.773
16045256 | 17859 | 4 | 2905E03271 | 2022-09-22T11:07:55.41+00:00 | FALSE | 4 | 79 | . | 152 | 24.5 | 2702.035
16045256 | 17859 | 1 | 2905E04011 | 2022-09-22T07:40:00.2+00:00 | TRUE | 1 | 0 | . | 152 | 24.5 | 2943.015
16045256 | 17859 | 3 | 2905E04833 | 2022-09-21T11:07:53.7+00:00 | TRUE | 2 | 0 | . | 152 | 24.5 | 3056.168
16045256 | 17859 | 2 | 2905E05086 | 2022-09-22T06:30:13.333+00:00 | TRUE | 3 | 0 | . | 152 | 24.5 | 3166.11
16045256 | 17859 | 2 | 2905E03507 | 2022-09-22T03:38:06.017+00:00 | TRUE | 3 | 0 | . | 152 | 24.5 | 3303.659
16045256 | 17859 | 2 | 2905E05074 | 2022-09-22T03:42:02.867+00:00 | FALSE | 3 | 46 | . | 152 | 24.5078 | 0
16045256 | 17859 | 1 | 2905E02667 | 2022-09-22T06:28:33.24+00:00 | TRUE | 1 | 0 | . | 152 | 24.5078 | 3028.286
16045256 | 17859 | 1 | 2905E05259 | 2022-09-22T07:43:54.867+00:00 | TRUE | 1 | 0 | . | 152 | 24.5078 | 3034.321
16045256 | 17859 | 1 | 2905E04358 | 2022-09-22T08:04:13.117+00:00 | TRUE | 1 | 0 | . | 152 | 24.5078 | 3044.435
16045256 | 17859 | 1 | 2905E03853 | 2022-09-22T07:49:18.487+00:00 | TRUE | 1 | 0 | . | 152 | 24.5078 | 3066.719
16045256 | 17859 | 1 | 2905E04743 | 2022-09-22T08:57:48.923+00:00 | TRUE | 1 | 0 | . | 152 | 24.5078 | 3073.208
16045256 | 17859 | 1 | 2905E04529 | 2022-09-22T06:31:02.943+00:00 | TRUE | 1 | 0 | . | 152 | 24.5078 | 3084.646
16045256 | 17859 | 2 | 2905E03665 | 2022-09-22T03:35:22.363+00:00 | TRUE | 3 | 0 | . | 152 | 24.5078 | 3227.026
16045256 | 17859 | 4 | 2905E04836 | 2022-09-22T09:49:24.01+00:00 | TRUE | 4 | 0 | . | 152 | 24.5156 | 2459.757
16045256 | 17859 | 4 | 2905E04982 | 2022-09-22T06:41:16.123+00:00 | TRUE | 4 | 0 | . | 152 | 24.5156 | 2526.809
16045256 | 17859 | 2 | 2905E04619 | 2022-09-22T10:22:34.967+00:00 | TRUE | 3 | 0 | . | 152 | 24.5156 | 3275.886
16045256 | 17859 | 4 | 2905E04688 | 2022-09-22T03:22:18.72+00:00 | FALSE | 4 | 88 | . | 152 | 24.5234 | -999
16045256 | 17859 | 4 | 2905E04323 | 2022-09-22T08:06:18.867+00:00 | TRUE | 4 | 0 | . | 152 | 24.5234 | 2460.024
16045256 | 17859 | 4 | 2905E04139 | 2022-09-22T10:52:25.797+00:00 | FALSE | 4 | 10 | . | 152 | 24.5234 | 2520.152
16045256 | 17859 | 1 | 2905E03744 | 2022-09-22T10:40:07.723+00:00 | TRUE | 1 | 0 | . | 152 | 24.5234 | 2969.768
16045256 | 17859 | 2 | 2905E04953 | 2022-09-22T02:44:42.7+00:00 | TRUE | 3 | 0 | . | 152 | 24.5234 | 3233.463
16045256 | 17859 | 4 | 2905E04564 | 2022-09-22T07:57:32.783+00:00 | FALSE | 4 | 88 | . | 152 | 24.5312 | -999
16045256 | 17859 | 2 | 2905E04947 | 2022-09-22T03:43:50.707+00:00 | FALSE | 3 | 88 | . | 152 | 24.5312 | -999
16045256 | 17859 | 4 | 2905E04730 | 2022-09-22T06:23:32.767+00:00 | TRUE | 4 | 0 | . | 152 | 24.5312 | 2459.659
16045256 | 17859 | 1 | 2905E05119 | 2022-09-22T02:33:40.78+00:00 | TRUE | 1 | 0 | . | 152 | 24.5312 | 2949.359
16045256 | 17859 | 3 | 2905E05277 | 2022-09-22T10:39:01.473+00:00 | TRUE | 2 | 0 | . | 152 | 24.5312 | 3034.591
16045256 | 17859 | 1 | 2905E04229 | 2022-09-22T08:16:42.163+00:00 | TRUE | 1 | 0 | . | 152 | 24.5312 | 3083.348
16045256 | 17859 | 1 | 2905E04837 | 2022-09-22T07:54:15.333+00:00 | TRUE | 1 | 0 | . | 152 | 24.5312 | 3091.154
16045256 | 17859 | 3 | 2905E04458 | 2022-09-22T03:22:00.097+00:00 | TRUE | 2 | 0 | . | 152 | 24.5312 | 3117.49
16045256 | 17859 | 1 | 2905E04134 | 2022-09-22T02:36:13.887+00:00 | TRUE | 1 | 0 | . | 152 | 24.5312 | 3124.529
16045256 | 17859 | 2 | 2905E04034 | 2022-09-22T06:24:15.257+00:00 | TRUE | 3 | 0 | . | 152 | 24.5312 | 3206.978
16045256 | 17859 | 2 | 2905E05018 | 2022-09-22T07:43:14.197+00:00 | TRUE | 3 | 0 | . | 152 | 24.5312 | 3252.941
16045256 | 17859 | 1 | 2905E05019 | 2022-09-22T08:38:31.453+00:00 | TRUE | 1 | 0 | . | 152 | 24.5391 | 3136.222
16045256 | 17859 | 4 | 2905E04748 | 2022-09-22T04:02:29.917+00:00 | TRUE | 4 | 0 | . | 152 | 24.5469 | 2486.422
16045256 | 17859 | 3 | 2905E04557 | 2022-09-22T03:12:22.547+00:00 | TRUE | 2 | 0 | . | 152 | 24.5469 | 2933.353
16045256 | 17859 | 1 | 2905E04553 | 2022-09-22T08:35:52.807+00:00 | TRUE | 1 | 0 | . | 152 | 24.5469 | 2988.113
16045256 | 17859 | 2 | 2905E05196 | 2022-09-22T07:33:02.39+00:00 | TRUE | 3 | 0 | . | 152 | 24.5469 | 3051.571
16045256 | 17859 | 2 | 2905E04286 | 2022-09-22T03:25:06.243+00:00 | TRUE | 3 | 0 | . | 152 | 24.5469 | 3196.289
16045256 | 17859 | 3 | 2905E05146 | 2022-09-22T10:25:23.3+00:00 | TRUE | 2 | 0 | . | 152 | 24.5469 | 3205.496
16045256 | 17859 | 3 | 2905E04664 | 2022-09-22T10:20:47.37+00:00 | FALSE | 2 | 88 | . | 152 | 24.5547 | -999
16045256 | 17859 | 1 | 2905E04959 | 2022-09-22T10:14:15.687+00:00 | FALSE | 1 | 88 | . | 152 | 24.5547 | -999
16045256 | 17859 | 1 | 2905E05013 | 2022-09-22T02:46:35.057+00:00 | TRUE | 1 | 0 | . | 152 | 24.5547 | 2957.548
16045256 | 17859 | 1 | 2905E04700 | 2022-09-22T03:23:11.453+00:00 | TRUE | 1 | 0 | . | 152 | 24.5547 | 3165.86
16045256 | 17859 | 2 | 2905E04939 | 2022-09-22T04:01:16.983+00:00 | TRUE | 3 | 0 | . | 152 | 24.5547 | 3265.7
16045256 | 17859 | 2 | 2905E04106 | 2022-09-22T07:55:58.77+00:00 | TRUE | 3 | 0 | . | 152 | 24.5547 | 3274.761
16045256 | 17859 | 4 | 2905E04354 | 2022-09-22T06:49:28.107+00:00 | TRUE | 4 | 0 | . | 152 | 24.5625 | 2403.277
16045256 | 17859 | 4 | 2905E04734 | 2022-09-22T03:53:35.777+00:00 | TRUE | 4 | 0 | . | 152 | 24.5625 | 2488.148
16045256 | 17859 | 4 | 2905E04534 | 2022-09-22T03:45:44.737+00:00 | TRUE | 4 | 0 | . | 152 | 24.5625 | 2521.375
16045256 | 17859 | 3 | 2905E04164 | 2022-09-22T03:00:40.553+00:00 | TRUE | 2 | 0 | . | 152 | 24.5625 | 2952.858
16045256 | 17859 | 2 | 2905E04016 | 2022-09-22T10:12:12.577+00:00 | TRUE | 3 | 0 | . | 152 | 24.5625 | 3237.154
16045256 | 17859 | 2 | 2905E05277 | 2022-09-22T10:13:46.037+00:00 | FALSE | 3 | 88 | . | 152 | 24.5703 | -999
16045256 | 17859 | 2 | 2905E04751 | 2022-09-22T07:16:56.317+00:00 | FALSE | 3 | 88 | . | 152 | 24.5703 | -999
16045256 | 17859 | 1 | 2905E04107 | 2022-09-22T08:33:25.43+00:00 | TRUE | 1 | 0 | . | 152 | 24.5703 | 3049.417
16045256 | 17859 | 2 | 2905E05040 | 2022-09-22T07:14:35.053+00:00 | TRUE | 3 | 0 | . | 152 | 24.5703 | 3166.105
16045256 | 17859 | 3 | 2905E04820 | 2022-09-22T03:14:54.947+00:00 | TRUE | 2 | 0 | . | 152 | 24.5781 | 2967.242
16045256 | 17859 | 1 | 2905E04749 | 2022-09-22T06:11:04.127+00:00 | TRUE | 1 | 0 | . | 152 | 24.5781 | 3081.062
16045256 | 17859 | 1 | 2905E04708 | 2022-09-22T08:14:21.173+00:00 | TRUE | 1 | 0 | . | 152 | 24.5781 | 3097.527
16045256 | 17859 | 2 | 2905E04594 | 2022-09-22T10:55:15.22+00:00 | FALSE | 3 | 69 | . | 152 | 24.5781 | 3169.706
16045256 | 17859 | 2 | 2905E04705 | 2022-09-22T03:40:28.527+00:00 | TRUE | 3 | 0 | . | 152 | 24.5781 | 3214.336
16045256 | 17859 | 2 | 2905E05171 | 2022-09-22T11:17:33.207+00:00 | TRUE | 3 | 0 | . | 152 | 24.5781 | 3282.871
16045256 | 17859 | 1 | 2905E03132 | 2022-09-22T08:44:02.677+00:00 | FALSE | 1 | 88 | . | 152 | 24.5859 | -999
16045256 | 17859 | 3 | 2905E04782 | 2022-09-22T10:26:59.37+00:00 | FALSE | 2 | 88 | . | 152 | 24.5938 | -999
16045256 | 17859 | 4 | 2905E03343 | 2022-09-22T03:51:04.5+00:00 | TRUE | 4 | 0 | . | 152 | 24.5938 | 2478.44
16045256 | 17859 | 2 | 2905E05068 | 2022-09-22T10:18:50.237+00:00 | TRUE | 3 | 0 | . | 152 | 24.5938 | 3099.901
16045256 | 17859 | 2 | 2905E01152 | 2022-09-22T03:04:10.617+00:00 | TRUE | 3 | 0 | . | 152 | 24.5938 | 3140.059
16045256 | 17859 | 2 | 2905E04348 | 2022-09-22T03:19:50.843+00:00 | TRUE | 3 | 0 | . | 152 | 24.5938 | 3196.469
16045256 | 17859 | 2 | 2905E04239 | 2022-09-22T08:31:10.96+00:00 | TRUE | 3 | 0 | . | 152 | 24.5938 | 3312.686
16045256 | 17859 | 1 | 2905E03946 | 2022-09-22T08:27:47.08+00:00 | FALSE | 1 | 88 | . | 152 | 24.6016 | -999
16045256 | 17859 | 4 | 2905E04354 | 2022-09-22T04:22:23.693+00:00 | FALSE | 4 | 88 | . | 152 | 24.6016 | -999
16045256 | 17859 | 4 | 2905E04020 | 2022-09-22T06:57:24.773+00:00 | TRUE | 4 | 0 | . | 152 | 24.6016 | 2483.53
16045256 | 17859 | 4 | 2905E03872 | 2022-09-22T07:51:10.087+00:00 | TRUE | 4 | 0 | . | 152 | 24.6016 | 2515.469
16045256 | 17859 | 1 | 2905E04742 | 2022-09-22T02:30:28.04+00:00 | TRUE | 1 | 0 | . | 152 | 24.6016 | 3001.833
16045256 | 17859 | 2 | 2905E03784 | 2022-09-22T03:47:32.597+00:00 | TRUE | 3 | 0 | . | 152 | 24.6016 | 3251.577
16045256 | 17859 | 2 | 2905E04736 | 2022-09-22T07:24:37.787+00:00 | FALSE | 3 | 88 | . | 152 | 24.6094 | -999
16045256 | 17859 | 2 | 2905E04992 | 2022-09-22T03:14:23.71+00:00 | FALSE | 3 | 88 | . | 152 | 24.6094 | -999
16045256 | 17859 | 4 | 2905E04327 | 2022-09-22T04:03:58.613+00:00 | FALSE | 4 | 34 | . | 152 | 24.6094 | 2501.819
16045256 | 17859 | 2 | 2905E05120 | 2022-09-22T11:07:42.027+00:00 | TRUE | 3 | 0 | . | 152 | 24.6094 | 3049.533
16045256 | 17859 | 2 | 2905E03884 | 2022-09-22T03:12:30.457+00:00 | TRUE | 3 | 0 | . | 152 | 24.6094 | 3109.772
16045256 | 17859 | 2 | 2905E05051 | 2022-09-22T03:58:27.05+00:00 | FALSE | 3 | 88 | . | 152 | 24.6172 | -999
16045256 | 17859 | 4 | 2905E03751 | 2022-09-22T08:41:37.057+00:00 | TRUE | 4 | 0 | . | 152 | 24.6172 | 2465.898
16045256 | 17859 | 1 | 2905E04384 | 2022-09-22T06:21:23.973+00:00 | TRUE | 1 | 0 | . | 152 | 24.6172 | 2935.014
16045256 | 17859 | 3 | 2905E04380 | 2022-09-22T03:19:24.347+00:00 | TRUE | 2 | 0 | . | 152 | 24.6172 | 3022.234
16045256 | 17859 | 4 | 2905E05127 | 2022-09-22T03:43:06.357+00:00 | FALSE | 4 | 88 | . | 152 | 24.625 | -999
16045256 | 17859 | 4 | 2905E03923 | 2022-09-22T08:26:27.88+00:00 | TRUE | 4 | 0 | . | 152 | 24.625 | 2576.323
16045256 | 17859 | 4 | 2905E05105 | 2022-09-22T10:32:24.86+00:00 | TRUE | 4 | 0 | . | 152 | 24.625 | 2587.728
16045256 | 17859 | 2 | 2905E04540 | 2022-09-22T10:06:35.733+00:00 | TRUE | 3 | 0 | . | 152 | 24.625 | 3055.816
16045256 | 17859 | 2 | 2905E04245 | 2022-09-22T08:26:26.94+00:00 | TRUE | 3 | 0 | . | 152 | 24.625 | 3255.845
16045256 | 17859 | 2 | 2905E03698 | 2022-09-22T03:54:42.543+00:00 | TRUE | 3 | 0 | . | 152 | 24.625 | 3279.622
16045256 | 17859 | 4 | 2905E04872 | 2022-09-22T07:00:44.147+00:00 | FALSE | 4 | 88 | . | 152 | 24.6328 | -999
16045256 | 17859 | 4 | 2905E04434 | 2022-09-22T08:39:04.16+00:00 | TRUE | 4 | 0 | . | 152 | 24.6328 | 2524.131
16045256 | 17859 | 4 | 2905E04429 | 2022-09-22T10:24:15.75+00:00 | FALSE | 4 | 77 | . | 152 | 24.6328 | 2538.337
16045256 | 17859 | 3 | 2905E04664 | 2022-09-22T10:32:36.38+00:00 | TRUE | 2 | 0 | . | 152 | 24.6328 | 2884.491
16045256 | 17859 | 3 | 2905E04445 | 2022-09-22T02:58:05.657+00:00 | TRUE | 2 | 0 | . | 152 | 24.6328 | 2950.271
16045256 | 17859 | 2 | 2905E03976 | 2022-09-22T04:29:29.61+00:00 | TRUE | 3 | 0 | . | 152 | 24.6328 | 3118.925
16045256 | 17859 | 2 | 2905E04702 | 2022-09-22T10:43:26.827+00:00 | TRUE | 3 | 0 | . | 152 | 24.6328 | 3289.976
16045256 | 17859 | 2 | 2905E03915 | 2022-09-22T08:02:27.997+00:00 | TRUE | 3 | 0 | . | 152 | 24.6328 | 3295.026
16045256 | 17859 | 2 | 2905E04593 | 2022-09-22T04:25:58.68+00:00 | TRUE | 3 | 0 | . | 152 | 24.6406 | 2989.99
16045256 | 17859 | 2 | 2905E04188 | 2022-09-22T08:52:08.033+00:00 | TRUE | 3 | 0 | . | 152 | 24.6406 | 3112.339
16045256 | 17859 | 1 | 2905E04290 | 2022-09-22T10:37:29.543+00:00 | TRUE | 1 | 0 | . | 152 | 24.6406 | 3113.052
16045256 | 17859 | 2 | 2905E05135 | 2022-09-22T03:22:21.943+00:00 | TRUE | 3 | 0 | . | 152 | 24.6406 | 3165.537
16045256 | 17859 | 2 | 2905E05217 | 2022-09-22T04:20:00.757+00:00 | TRUE | 3 | 0 | . | 152 | 24.6406 | 3177.91
16045256 | 17859 | 2 | 2905E04679 | 2022-09-22T04:17:19.3+00:00 | TRUE | 3 | 0 | . | 152 | 24.6406 | 3274.856
16045256 | 17859 | 1 | 2905E05120 | 2022-09-22T10:42:03.4+00:00 | FALSE | 1 | 88 | . | 152 | 24.6484 | -999
16045256 | 17859 | 2 | 2905E03271 | 2022-09-22T10:15:29.967+00:00 | FALSE | 3 | 88 | . | 152 | 24.6484 | -999
16045256 | 17859 | 1 | 2905E03705 | 2022-09-22T03:19:08.423+00:00 | FALSE | 1 | 88 | . | 152 | 24.6484 | -999
16045256 | 17859 | 3 | 2905E04028 | 2022-09-22T03:16:41.42+00:00 | FALSE | 2 | 88 | . | 152 | 24.6484 | -999
16045256 | 17859 | 4 | 2905E04310 | 2022-09-22T03:59:21.433+00:00 | FALSE | 4 | 70 | . | 152 | 24.6484 | 2548.847
16045256 | 17859 | 1 | 2905E03903 | 2022-09-22T03:18:01.27+00:00 | TRUE | 1 | 0 | . | 152 | 24.6484 | 3107.211
16045256 | 17859 | 1 | 2905E04827 | 2022-09-22T08:55:11.233+00:00 | TRUE | 1 | 0 | . | 152 | 24.6484 | 3143.843
16045256 | 17859 | 1 | 2905E04987 | 2022-09-22T08:30:40.48+00:00 | FALSE | 1 | 88 | . | 152 | 24.6562 | -999
16045256 | 17859 | 4 | 2905E04543 | 2022-09-22T08:28:32.693+00:00 | FALSE | 4 | 88 | . | 152 | 24.6562 | -999
16045256 | 17859 | 3 | 2905E03820 | 2022-09-22T03:05:15.9+00:00 | FALSE | 2 | 88 | . | 152 | 24.6562 | -999
16045256 | 17859 | 1 | 2905E04594 | 2022-09-22T06:40:40.2+00:00 | FALSE | 1 | 80 | . | 152 | 24.6641 | 3004.202
16045256 | 17859 | 2 | 2905E05128 | 2022-09-22T08:11:32.967+00:00 | TRUE | 3 | 0 | . | 152 | 24.6641 | 3188.979
16045256 | 17859 | 2 | 2905E04738 | 2022-09-22T06:35:59.407+00:00 | TRUE | 3 | 0 | . | 152 | 24.6641 | 3212.858
16045256 | 17859 | 2 | 2905E03833 | 2022-09-22T08:17:02.033+00:00 | TRUE | 3 | 0 | . | 152 | 24.6641 | 3278.316
16045256 | 17859 | 1 | 2905E04119 | 2022-09-22T10:28:39.89+00:00 | FALSE | 1 | 88 | . | 152 | 24.6719 | -999
16045256 | 17859 | 1 | 2905E04217 | 2022-09-22T08:52:14.2+00:00 | FALSE | 1 | 88 | . | 152 | 24.6719 | -999
16045256 | 17859 | 1 | 2905E04032 | 2022-09-22T03:25:44.933+00:00 | TRUE | 1 | 0 | . | 152 | 24.6719 | 3038.435
16045256 | 17859 | 2 | 2905E04006 | 2022-09-22T10:27:36.247+00:00 | TRUE | 3 | 0 | . | 152 | 24.6719 | 3286.282
16045256 | 17859 | 1 | 2905E04632 | 2022-09-22T11:09:29.263+00:00 | FALSE | 1 | 88 | . | 152 | 24.6797 | -999
16045256 | 17859 | 1 | 2905E04730 | 2022-09-22T06:03:56.95+00:00 | FALSE | 1 | 88 | . | 152 | 24.6797 | -999
16045256 | 17859 | 1 | 2905E03631 | 2022-09-22T08:11:39.703+00:00 | TRUE | 1 | 0 | . | 152 | 24.6797 | 3053.773
16045256 | 17859 | 2 | 2905E03180 | 2022-09-22T04:14:53.68+00:00 | TRUE | 3 | 0 | . | 152 | 24.6797 | 3156.266
16045256 | 17859 | 4 | 2905E04669 | 2022-09-22T04:12:03.037+00:00 | TRUE | 4 | 0 | . | 152 | 24.6875 | 2473.329
16045256 | 17859 | 1 | 2905E04547 | 2022-09-22T06:54:02.917+00:00 | TRUE | 1 | 0 | . | 152 | 24.6875 | 3008.239
16045256 | 17859 | 2 | 2905E04021 | 2022-09-22T04:08:48.65+00:00 | TRUE | 3 | 0 | . | 152 | 24.6875 | 3127.924
16045256 | 17859 | 2 | 2905E03767 | 2022-09-22T08:42:27.997+00:00 | TRUE | 3 | 0 | . | 152 | 24.6875 | 3268.411
16045256 | 17859 | 1 | 2905E04662 | 2022-09-22T10:32:10.563+00:00 | FALSE | 1 | 88 | . | 152 | 24.6953 | -999
16045256 | 17859 | 4 | 2905E05040 | 2022-09-22T06:58:50.433+00:00 | FALSE | 4 | 73 | . | 152 | 24.6953 | 2455.718
16045256 | 17859 | 4 | 2905E04376 | 2022-09-22T08:36:37.237+00:00 | TRUE | 4 | 0 | . | 152 | 24.6953 | 2540.669
16045256 | 17859 | 2 | 2905E04158 | 2022-09-22T07:53:33.25+00:00 | TRUE | 3 | 0 | . | 152 | 24.6953 | 3108.225
16045256 | 17859 | 1 | 2905E04873 | 2022-09-22T08:59:25.77+00:00 | FALSE | 1 | 88 | . | 152 | 24.7031 | -999
16045256 | 17859 | 2 | 2905E04347 | 2022-09-22T08:54:54.86+00:00 | TRUE | 3 | 0 | . | 152 | 24.7031 | 3148.781
16045256 | 17859 | 2 | 2905E04811 | 2022-09-22T11:11:47.013+00:00 | FALSE | 3 | 88 | . | 152 | 24.7109 | -999
16045256 | 17859 | 1 | 2905E04298 | 2022-09-22T06:16:21.213+00:00 | TRUE | 1 | 0 | . | 152 | 24.7109 | 2979.436
16045256 | 17859 | 1 | 2905E04959 | 2022-09-22T11:13:39.523+00:00 | TRUE | 1 | 0 | . | 152 | 24.7109 | 3150.707
16045256 | 17859 | 2 | 2905E04295 | 2022-09-22T08:23:01.933+00:00 | TRUE | 3 | 0 | . | 152 | 24.7109 | 3188.042
16045256 | 17859 | 1 | 2905E03940 | 2022-09-22T07:56:29.03+00:00 | FALSE | 1 | 98 | . | 152 | 24.7109 | 3191.698
16045256 | 17859 | 4 | 2905E03946 | 2022-09-22T08:31:26.963+00:00 | FALSE | 4 | 88 | . | 152 | 24.7188 | -999
16045256 | 17859 | 1 | 2905E04654 | 2022-09-22T02:57:12.31+00:00 | TRUE | 1 | 0 | . | 152 | 24.7188 | 2939.538
16045256 | 17859 | 1 | 2905E04435 | 2022-09-22T08:47:01.437+00:00 | TRUE | 1 | 0 | . | 152 | 24.7188 | 3021.269
16045256 | 17859 | 1 | 2905E04952 | 2022-09-22T08:20:05.42+00:00 | TRUE | 1 | 0 | . | 152 | 24.7188 | 3074.331
16045256 | 17859 | 2 | 2905E04383 | 2022-09-22T08:20:10.927+00:00 | TRUE | 3 | 0 | . | 152 | 24.7188 | 3307.213
16045256 | 17859 | 2 | 2905E03746 | 2022-09-22T10:46:18.577+00:00 | TRUE | 3 | 0 | . | 152 | 24.7188 | 3380.702
16045256 | 17859 | 2 | 2905E03693 | 2022-09-22T03:49:28.28+00:00 | FALSE | 3 | 88 | . | 152 | 24.7266 | -999
16045256 | 17859 | 1 | 2905E04012 | 2022-09-22T06:48:56.377+00:00 | FALSE | 1 | 68 | . | 152 | 24.7266 | 2935.064
16045256 | 17859 | 1 | 2905E03925 | 2022-09-22T03:15:39.593+00:00 | TRUE | 1 | 0 | . | 152 | 24.7266 | 3045.233
16045256 | 17859 | 1 | 2905E05126 | 2022-09-22T02:49:33.623+00:00 | TRUE | 1 | 0 | . | 152 | 24.7266 | 3046.033
16045256 | 17859 | 1 | 2905E04958 | 2022-09-22T06:44:28.483+00:00 | FALSE | 1 | 81 | . | 152 | 24.7266 | 3051.288
16045256 | 17859 | 1 | 2905E04258 | 2022-09-22T11:02:10.087+00:00 | TRUE | 1 | 0 | . | 152 | 24.7266 | 3076.191
16045256 | 17859 | 2 | 2905E04378 | 2022-09-22T04:23:06.037+00:00 | TRUE | 3 | 0 | . | 152 | 24.7266 | 3082.645
16045256 | 17859 | 2 | 2905E04310 | 2022-09-22T06:39:16.13+00:00 | TRUE | 3 | 0 | . | 152 | 24.7266 | 3174.422
16045256 | 17859 | 2 | 2905E03808 | 2022-09-22T08:14:21.733+00:00 | TRUE | 3 | 0 | . | 152 | 24.7266 | 3258.86
16045256 | 17859 | 2 | 2905E04946 | 2022-09-22T03:56:19.937+00:00 | FALSE | 3 | 88 | . | 152 | 24.7344 | -999
16045256 | 17859 | 3 | 2905E04320 | 2022-09-22T03:03:21.827+00:00 | TRUE | 2 | 0 | . | 152 | 24.7344 | 3009.024
16045256 | 17859 | 1 | 2905E04448 | 2022-09-22T03:12:52.373+00:00 | TRUE | 1 | 0 | . | 152 | 24.7344 | 3090.68
16045256 | 17859 | 1 | 2905E04540 | 2022-09-22T06:37:56.317+00:00 | FALSE | 1 | 69 | . | 152 | 24.7422 | 2955.919
16045256 | 17859 | 1 | 2905E04521 | 2022-09-22T06:51:39.397+00:00 | TRUE | 1 | 0 | . | 152 | 24.7422 | 2977.012
16045256 | 17859 | 1 | 2905E04429 | 2022-09-22T10:34:49.037+00:00 | TRUE | 1 | 0 | . | 152 | 24.7422 | 3056.097
16045256 | 17859 | 2 | 2905E04959 | 2022-09-22T10:40:12.613+00:00 | FALSE | 3 | 49 | . | 152 | 24.75 | 0
16045256 | 17859 | 1 | 2905E04957 | 2022-09-22T06:25:59.57+00:00 | FALSE | 1 | 46 | . | 152 | 24.75 | 0
16045256 | 17859 | 4 | 2905E03279 | 2022-09-22T08:52:08.173+00:00 | TRUE | 4 | 0 | . | 152 | 24.75 | 2432.338
16045256 | 17859 | 2 | 2905E04601 | 2022-09-22T07:58:38.293+00:00 | TRUE | 3 | 0 | . | 152 | 24.75 | 3260.358
16045256 | 17859 | 1 | 2905E04628 | 2022-09-22T02:54:45.973+00:00 | TRUE | 1 | 0 | . | 152 | 24.7578 | 2939.839
16045256 | 17859 | 3 | 2905E03811 | 2022-09-22T03:09:38.957+00:00 | FALSE | 2 | 81 | . | 152 | 24.7578 | 3017.663
16045256 | 17859 | 3 | 2905E04896 | 2022-09-22T03:54:48.317+00:00 | TRUE | 2 | 0 | . | 152 | 24.7578 | 3099.034
16045256 | 17859 | 2 | 2905E04093 | 2022-09-22T03:09:30.73+00:00 | TRUE | 3 | 0 | . | 152 | 24.7578 | 3121.652
16045256 | 17859 | 2 | 2905E04035 | 2022-09-22T06:41:57.15+00:00 | TRUE | 3 | 0 | . | 152 | 24.7578 | 3127.179
16045256 | 17859 | 1 | 2905E04751 | 2022-09-22T10:47:50.217+00:00 | TRUE | 1 | 0 | . | 152 | 24.7578 | 3149.75
16045256 | 17859 | 2 | 2905E04776 | 2022-09-22T08:08:47.987+00:00 | TRUE | 3 | 0 | . | 152 | 24.7578 | 3199.522
16045256 | 17859 | 3 | 2905E03271 | 2022-09-22T10:36:05.863+00:00 | FALSE | 2 | 88 | . | 152 | 24.7656 | -999
16045256 | 17859 | 4 | 2905E04955 | 2022-09-22T08:44:22.56+00:00 | TRUE | 4 | 0 | . | 152 | 24.7656 | 2471.588
16045256 | 17859 | 1 | 2905E04012 | 2022-09-22T10:53:40.793+00:00 | FALSE | 1 | 68 | . | 152 | 24.7656 | 2937.581
16045256 | 17859 | 1 | 2905E04661 | 2022-09-22T06:24:16.067+00:00 | FALSE | 1 | 60 | . | 152 | 24.7734 | 3011.241
16045256 | 17859 | 3 | 2905E04732 | 2022-09-22T03:32:30.89+00:00 | TRUE | 2 | 0 | . | 152 | 24.7734 | 3052.792
16045256 | 17859 | 2 | 2905E04575 | 2022-09-22T04:04:24.493+00:00 | TRUE | 3 | 0 | . | 152 | 24.7734 | 3128.189
16045256 | 17859 | 2 | 2905E03926 | 2022-09-22T10:52:24.49+00:00 | TRUE | 3 | 0 | . | 152 | 24.7734 | 3265.694
16045256 | 17859 | 2 | 2905E03982 | 2022-09-22T10:30:19.077+00:00 | TRUE | 3 | 0 | . | 152 | 24.7734 | 3325.287
16045256 | 17859 | 2 | 2905E05012 | 2022-09-22T11:15:06.453+00:00 | TRUE | 3 | 0 | . | 152 | 24.7812 | 3054.209
16045256 | 17859 | 1 | 2905E04321 | 2022-09-22T01:54:32.97+00:00 | TRUE | 1 | 0 | . | 152 | 24.7812 | 3091.88
16045256 | 17859 | 1 | 2905E04814 | 2022-09-22T03:33:38.263+00:00 | TRUE | 1 | 0 | . | 152 | 24.7812 | 3120.514
16045256 | 17859 | 2 | 2905E04250 | 2022-09-22T04:11:53.637+00:00 | TRUE | 3 | 0 | . | 152 | 24.7891 | 2978.415
16045256 | 17859 | 3 | 2905E04012 | 2022-09-22T04:08:43.443+00:00 | FALSE | 2 | 70 | . | 152 | 24.7969 | 2890.484
16045256 | 17859 | 3 | 2905E03851 | 2022-09-22T04:03:47.087+00:00 | FALSE | 2 | 34 | . | 152 | 24.7969 | 2986.076
16045256 | 17859 | 2 | 2905E03818 | 2022-09-22T03:52:13.62+00:00 | TRUE | 3 | 0 | . | 152 | 24.7969 | 3132.703
16045256 | 17859 | 2 | 2905E01216 | 2022-09-22T11:01:02.35+00:00 | TRUE | 3 | 0 | . | 152 | 24.7969 | 3169.646
16045256 | 17859 | 3 | 2905E04322 | 2022-09-22T03:59:10.84+00:00 | TRUE | 2 | 0 | . | 152 | 24.8047 | 3033.46
16045256 | 17859 | 3 | 2905E05155 | 2022-09-22T03:40:28.777+00:00 | TRUE | 2 | 0 | . | 152 | 24.8047 | 3034.892
16045256 | 17859 | 1 | 2905E03763 | 2022-09-22T10:59:46.343+00:00 | TRUE | 1 | 0 | . | 152 | 24.8047 | 3115.973
16045256 | 17859 | 2 | 2905E04739 | 2022-09-22T07:48:26.693+00:00 | TRUE | 3 | 0 | . | 152 | 24.8047 | 3164.58
16045256 | 17859 | 3 | 2905E04932 | 2022-09-22T03:35:15.94+00:00 | TRUE | 2 | 0 | . | 152 | 24.8125 | 3020.711
16045256 | 17859 | 3 | 2905E04484 | 2022-09-22T03:27:23.58+00:00 | TRUE | 2 | 0 | . | 152 | 24.8125 | 3032.001
16045256 | 17859 | 4 | 2905E05012 | 2022-09-22T08:48:57.397+00:00 | FALSE | 4 | 88 | . | 152 | 24.8203 | -999
16045256 | 17859 | 2 | 2905E05012 | 2022-09-22T08:43:59.657+00:00 | FALSE | 3 | 88 | . | 152 | 24.8203 | -999
16045256 | 17859 | 2 | 2905E03980 | 2022-09-22T10:49:25.287+00:00 | TRUE | 3 | 0 | . | 152 | 24.8203 | 3337.981
16045256 | 17859 | 3 | 2905E05336 | 2022-09-22T03:28:46.073+00:00 | FALSE | 2 | 81 | . | 152 | 24.8281 | 2894.714
16045256 | 17859 | 2 | 2905E04217 | 2022-09-22T10:58:16.983+00:00 | TRUE | 3 | 0 | . | 152 | 24.8281 | 3213.833
16045256 | 17859 | 2 | 2905E04258 | 2022-09-22T08:48:31.307+00:00 | FALSE | 3 | 88 | . | 152 | 24.8438 | -999
16045256 | 17859 | 2 | 2905E04837 | 2022-09-22T07:50:04.603+00:00 | FALSE | 3 | 88 | . | 152 | 24.8438 | -999
16045256 | 17859 | 2 | 2905E04946 | 2022-09-22T06:52:01.413+00:00 | TRUE | 3 | 0 | . | 152 | 24.8438 | 3118.29
16045256 | 17859 | 2 | 2905E04662 | 2022-09-22T11:10:08.153+00:00 | TRUE | 3 | 0 | . | 152 | 24.8438 | 3295.092
16045256 | 17859 | 1 | 2905E04242 | 2022-09-22T06:55:42.747+00:00 | FALSE | 1 | 88 | . | 152 | 24.8516 | -999
16045256 | 17859 | 1 | 2905E04535 | 2022-09-22T03:20:16.25+00:00 | FALSE | 1 | 46 | . | 152 | 24.8516 | 0
16045256 | 17859 | 1 | 2905E04373 | 2022-09-22T03:02:12.947+00:00 | TRUE | 1 | 0 | . | 152 | 24.8516 | 2942.45
16045256 | 17859 | 1 | 2905E04242 | 2022-09-22T10:50:32.917+00:00 | TRUE | 1 | 0 | . | 152 | 24.8516 | 2977.533
16045256 | 17859 | 1 | 2905E03643 | 2022-09-22T04:09:45.23+00:00 | TRUE | 1 | 0 | . | 152 | 24.8516 | 2991.696
16045256 | 17859 | 3 | 2905E05078 | 2022-09-22T04:25:44.677+00:00 | TRUE | 2 | 0 | . | 152 | 24.8516 | 3056.977
16045256 | 17859 | 1 | 2905E04310 | 2022-09-22T03:05:04.07+00:00 | FALSE | 1 | 68 | . | 152 | 24.8516 | 3086.271
16045256 | 17859 | 3 | 2905E05078 | 2022-09-22T03:49:01.19+00:00 | FALSE | 2 | 59 | . | 152 | 24.8516 | 3094.927
16045256 | 17859 | 2 | 2905E04000 | 2022-09-22T06:49:08.217+00:00 | TRUE | 3 | 0 | . | 152 | 24.8516 | 3123.106
16045256 | 17859 | 4 | 2905E03132 | 2022-09-22T08:47:17.56+00:00 | TRUE | 4 | 0 | . | 152 | 24.8594 | 2532.624
16045256 | 17859 | 1 | 2905E03777 | 2022-09-22T04:16:52.347+00:00 | TRUE | 1 | 0 | . | 152 | 24.8594 | 2972.019
16045256 | 17859 | 3 | 2905E04948 | 2022-09-22T03:24:41.193+00:00 | TRUE | 2 | 0 | . | 152 | 24.8594 | 3024.26
16045256 | 17859 | 1 | 2905E05011 | 2022-09-22T03:42:27.847+00:00 | TRUE | 1 | 0 | . | 152 | 24.8594 | 3072.476
16045256 | 17859 | 2 | 2905E03617 | 2022-09-22T08:57:22.277+00:00 | TRUE | 3 | 0 | . | 152 | 24.8594 | 3151.431
16045256 | 17859 | 4 | 2905E04866 | 2022-09-22T11:17:30.777+00:00 | FALSE | 4 | 46 | . | 152 | 24.8672 | 0
16045256 | 17859 | 1 | 2905E04957 | 2022-09-22T03:39:50.897+00:00 | FALSE | 1 | 46 | . | 152 | 24.8672 | 0
16045256 | 17859 | 1 | 2905E03873 | 2022-09-22T02:59:45.287+00:00 | TRUE | 1 | 0 | . | 152 | 24.875 | 3011.834
16045256 | 17859 | 2 | 2905E05198 | 2022-09-22T09:00:21.167+00:00 | TRUE | 3 | 0 | . | 152 | 24.875 | 3224.609
16045256 | 17859 | 2 | 2905E04139 | 2022-09-22T06:55:16.697+00:00 | FALSE | 3 | 10 | . | 152 | 24.875 | 3260.626
16045256 | 17859 | 1 | 2905E04543 | 2022-09-22T08:24:52.833+00:00 | FALSE | 1 | 88 | . | 152 | 24.8828 | -999
16045256 | 17859 | 3 | 2905E04856 | 2022-09-22T03:52:17.433+00:00 | TRUE | 2 | 0 | . | 152 | 24.8906 | 3064.079
16045256 | 17859 | 3 | 2905E04906 | 2022-09-22T03:56:09.883+00:00 | FALSE | 2 | 88 | . | 152 | 24.8984 | -999
16045256 | 17859 | 2 | 2905E04728 | 2022-09-22T07:00:58.82+00:00 | TRUE | 3 | 0 | . | 152 | 24.8984 | 3154.051
16045256 | 17859 | 2 | 2905E04913 | 2022-09-22T10:33:00.987+00:00 | TRUE | 3 | 0 | . | 152 | 24.8984 | 3257.2
16045256 | 17859 | 1 | 2905E03762 | 2022-09-22T10:57:06.793+00:00 | TRUE | 1 | 0 | . | 152 | 24.9062 | 2976.359
16045256 | 17859 | 1 | 2905E03550 | 2022-09-22T03:55:29.153+00:00 | TRUE | 1 | 0 | . | 152 | 24.9062 | 2976.956
16045256 | 17859 | 1 | 2905E04200 | 2022-09-22T03:28:15.677+00:00 | TRUE | 1 | 0 | . | 152 | 24.9062 | 3075.194
16045256 | 17859 | 3 | 2905E04431 | 2022-09-22T03:37:42.167+00:00 | TRUE | 2 | 0 | . | 152 | 24.9062 | 3164.232
16045256 | 17859 | 1 | 2905E04736 | 2022-09-22T11:16:25.797+00:00 | TRUE | 1 | 0 | . | 152 | 24.9219 | 3003.679
16045256 | 17859 | 1 | 2905E04615 | 2022-09-22T03:58:12.367+00:00 | TRUE | 1 | 0 | . | 152 | 24.9297 | 2952.664
16045256 | 17859 | 2 | 2905E04670 | 2022-09-22T08:39:54.68+00:00 | TRUE | 3 | 0 | . | 152 | 24.9297 | 3180.86
16045256 | 17859 | 2 | 2905E04987 | 2022-09-22T08:33:33.447+00:00 | TRUE | 3 | 0 | . | 152 | 24.9297 | 3279.216
16045256 | 17859 | 3 | 2905E03847 | 2022-09-22T04:02:26.797+00:00 | TRUE | 2 | 0 | . | 152 | 24.9375 | 3040.215
16045256 | 17859 | 2 | 2905E04782 | 2022-09-22T11:03:10.417+00:00 | FALSE | 3 | 88 | . | 152 | 24.9453 | -999
16045256 | 17859 | 1 | 2905E04639 | 2022-09-22T03:10:09.26+00:00 | TRUE | 1 | 0 | . | 152 | 24.9453 | 2960.206
16045256 | 17859 | 1 | 2905E03135 | 2022-09-22T03:36:14.05+00:00 | TRUE | 1 | 0 | . | 152 | 24.9453 | 3000.374
16045256 | 17859 | 3 | 2905E04908 | 2022-09-22T04:20:12.543+00:00 | TRUE | 2 | 0 | . | 152 | 24.9453 | 3006.081
16045256 | 17859 | 1 | 2905E04469 | 2022-09-22T10:44:33.093+00:00 | FALSE | 1 | 88 | . | 152 | 24.9531 | -999
16045256 | 17859 | 3 | 2905E04507 | 2022-09-22T04:29:23.783+00:00 | TRUE | 2 | 0 | . | 152 | 24.9531 | 3034.99
16045256 | 17859 | 2 | 2905E04327 | 2022-09-22T06:46:19.95+00:00 | TRUE | 3 | 0 | . | 152 | 24.9609 | 3186.565
16045256 | 17859 | 1 | 2905E03438 | 2022-09-22T03:38:42.733+00:00 | TRUE | 1 | 0 | . | 152 | 24.9688 | 2920.5
16045256 | 17859 | 3 | 2905E04832 | 2022-09-22T03:44:08.62+00:00 | FALSE | 2 | 88 | . | 152 | 24.9766 | -999
16045256 | 17859 | 1 | 2905E05023 | 2022-09-22T03:44:49.8+00:00 | TRUE | 1 | 0 | . | 152 | 24.9766 | 2944.193
16045256 | 17859 | 2 | 2905E04782 | 2022-09-22T10:35:11.287+00:00 | FALSE | 3 | 88 | . | 152 | 24.9844 | -999
16045256 | 17859 | 1 | 2905E04119 | 2022-09-22T11:08:22.38+00:00 | TRUE | 1 | 0 | . | 152 | 24.9844 | 3026.206
16045256 | 17859 | 2 | 2905E03952 | 2022-09-22T10:38:24.117+00:00 | TRUE | 3 | 0 | . | 152 | 24.9844 | 3150.57
16045256 | 17859 | 3 | 2905E04457 | 2022-09-22T04:17:16.337+00:00 | TRUE | 2 | 0 | . | 152 | 24.9922 | 2938.616
16045256 | 17859 | 2 | 2905E04242 | 2022-09-22T06:58:23.257+00:00 | FALSE | 3 | 88 | . | 152 | 25 | -999
16045256 | 17859 | 1 | 2905E05051 | 2022-09-22T04:03:35.787+00:00 | TRUE | 1 | 0 | . | 152 | 25.0078 | 2941.448
16045256 | 17859 | 4 | 2905E04217 | 2022-09-22T08:58:36.677+00:00 | FALSE | 4 | 88 | . | 152 | 25.0156 | -999
16045256 | 17859 | 1 | 2905E04146 | 2022-09-22T03:53:06.677+00:00 | TRUE | 1 | 0 | . | 152 | 25.0156 | 3094.745
16045256 | 17859 | 1 | 2905E03896 | 2022-09-22T03:50:27.947+00:00 | TRUE | 1 | 0 | . | 152 | 25.0156 | 3134.405
16045256 | 17859 | 3 | 2905E03385 | 2022-09-22T04:23:00.597+00:00 | TRUE | 2 | 0 | . | 152 | 25.0312 | 2842.21
16045256 | 17859 | 1 | 2905E05017 | 2022-09-22T03:30:50.317+00:00 | TRUE | 1 | 0 | . | 152 | 25.0391 | 3057.179
16045256 | 17859 | 3 | 2905E02667 | 2022-09-22T03:42:10.897+00:00 | FALSE | 2 | 88 | . | 152 | 25.0547 | -999
16045256 | 17859 | 2 | 2905E03938 | 2022-09-22T08:05:03.713+00:00 | FALSE | 3 | 88 | . | 152 | 25.0625 | -999
16045256 | 17859 | 3 | 2905E04929 | 2022-09-22T03:47:32.427+00:00 | TRUE | 2 | 0 | . | 152 | 25.0625 | 3028.06
16045256 | 17859 | 1 | 2905E04139 | 2022-09-22T06:58:42.833+00:00 | FALSE | 1 | 10 | . | 152 | 25.0625 | 3073.34
16045256 | 17859 | 3 | 2905E04848 | 2022-09-22T04:14:30.417+00:00 | TRUE | 2 | 0 | . | 152 | 25.0781 | 3083.405
16045256 | 17859 | 1 | 2905E04690 | 2022-09-22T04:11:08.807+00:00 | FALSE | 1 | 88 | . | 152 | 25.0859 | -999
16045256 | 17859 | 1 | 2905E04906 | 2022-09-22T04:01:06.527+00:00 | TRUE | 1 | 0 | . | 152 | 25.0859 | 3033.059
16045256 | 17859 | 4 | 2905E04873 | 2022-09-22T09:01:19.38+00:00 | FALSE | 4 | 88 | . | 152 | 25.0938 | -999
16045256 | 17859 | 3 | 2905E03733 | 2022-09-22T04:11:35.737+00:00 | TRUE | 2 | 0 | . | 152 | 25.0938 | 2942.613
16045256 | 17859 | 1 | 2905E03180 | 2022-09-22T04:19:29.743+00:00 | TRUE | 1 | 0 | . | 152 | 25.1094 | 2893.751
16045256 | 17859 | 1 | 2905E04667 | 2022-09-22T04:21:52.277+00:00 | TRUE | 1 | 0 | . | 152 | 25.1094 | 3113.06
16045256 | 17859 | 1 | 2905E04878 | 2022-09-22T03:47:54.837+00:00 | TRUE | 1 | 0 | . | 152 | 25.125 | 3115.185
16045256 | 17859 | 1 | 2905E04459 | 2022-09-22T04:14:19.007+00:00 | TRUE | 1 | 0 | . | 152 | 25.1406 | 2967.394
16045256 | 17859 | 1 | 2905E04646 | 2022-09-22T04:27:12.033+00:00 | TRUE | 1 | 0 | . | 152 | 25.1406 | 2981.984
16045256 | 17859 | 1 | 2905E04677 | 2022-09-22T11:19:21.297+00:00 | TRUE | 1 | 0 | . | 152 | 25.2109 | 2989.875
16045256 | 17859 | 4 | 2905E04913 | 2022-09-22T10:14:24.077+00:00 | FALSE | 4 | 41 | . | 152 | 47.3047 | 0
16045256 | 17858 | 1 | 2905E04119 | 2022-09-22T10:23:51.407+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04597 | 2022-09-22T09:51:19.7+00:00 | FALSE | 2 | 20 | . | 704
16045256 | 17858 | 1 | 2905E04159 | 2022-09-22T09:48:25.323+00:00 | FALSE | 2 | 20 | . | 704
16045256 | 17858 | 1 | 2905E04543 | 2022-09-22T09:45:25.85+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04543 | 2022-09-22T09:43:04.4+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03774 | 2022-09-22T08:20:03.13+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 4 | 2905E05019 | 2022-09-22T08:17:40.703+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04471 | 2022-09-22T08:15:32.803+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04434 | 2022-09-22T08:12:38.23+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03876 | 2022-09-22T08:11:59.957+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03767 | 2022-09-22T08:10:09.267+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04258 | 2022-09-22T08:09:59.207+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05012 | 2022-09-22T08:09:47.36+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03751 | 2022-09-22T08:07:42.173+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03942 | 2022-09-22T08:07:24.393+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04670 | 2022-09-22T08:07:13.227+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04435 | 2022-09-22T08:05:24.093+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03132 | 2022-09-22T08:05:07.063+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04955 | 2022-09-22T08:04:55.703+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03279 | 2022-09-22T08:02:59.387+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04188 | 2022-09-22T08:02:46.197+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03762 | 2022-09-22T08:02:33.177+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04371 | 2022-09-22T07:59:51.813+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04827 | 2022-09-22T07:59:38.843+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03617 | 2022-09-22T07:59:30.837+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04347 | 2022-09-22T07:57:23.297+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04217 | 2022-09-22T07:57:13.613+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05114 | 2022-09-22T07:57:03.037+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04873 | 2022-09-22T07:54:56.637+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05198 | 2022-09-22T07:54:47.587+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04743 | 2022-09-22T07:54:38.827+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04866 | 2022-09-22T07:52:33.323+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05048 | 2022-09-22T07:52:19.523+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03961 | 2022-09-22T07:52:09.467+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03133 | 2022-09-22T07:50:20.237+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04578 | 2022-09-22T07:49:59.113+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05019 | 2022-09-22T07:49:46.29+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04836 | 2022-09-22T07:49:43.687+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05030 | 2022-09-22T07:46:41.887+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03641 | 2022-09-22T07:46:28.843+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04436 | 2022-09-22T07:46:18.5+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05161 | 2022-09-22T07:44:02.083+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04825 | 2022-09-22T07:43:52.103+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04676 | 2022-09-22T07:43:43.213+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05203 | 2022-09-22T07:41:37.82+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03700 | 2022-09-22T07:41:15.123+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04434 | 2022-09-22T07:41:04.343+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04290 | 2022-09-22T07:41:03.923+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03885 | 2022-09-22T07:39:12.113+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04013 | 2022-09-22T07:38:52.53+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04141 | 2022-09-22T07:38:40.973+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04471 | 2022-09-22T07:38:35.317+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04943 | 2022-09-22T07:36:25.533+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04271 | 2022-09-22T07:36:22.663+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03744 | 2022-09-22T07:36:07.923+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03922 | 2022-09-22T07:35:49.843+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05249 | 2022-09-22T07:33:59.303+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04913 | 2022-09-22T07:33:57.907+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03695 | 2022-09-22T07:33:14.37+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04704 | 2022-09-22T07:33:05.61+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05171 | 2022-09-22T07:31:42.81+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03888 | 2022-09-22T07:31:37.24+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04011 | 2022-09-22T07:30:30.067+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04677 | 2022-09-22T07:30:19.877+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03888 | 2022-09-22T07:29:19.123+00:00 | FALSE | 4 | 7 | . | 704
16045256 | 17858 | 2 | 2905E03579 | 2022-09-22T07:28:02.483+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04959 | 2022-09-22T07:27:46.72+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03952 | 2022-09-22T07:27:35.003+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05120 | 2022-09-22T07:27:25.057+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04077 | 2022-09-22T07:24:18.387+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04014 | 2022-09-22T07:24:15.917+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05277 | 2022-09-22T07:23:59.927+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03774 | 2022-09-22T07:23:26.643+00:00 | FALSE | 2 | 30 | . | 704
16045256 | 17858 | 2 | 2905E04438 | 2022-09-22T07:21:28.887+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04811 | 2022-09-22T07:21:12.217+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03774 | 2022-09-22T07:21:06.697+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04016 | 2022-09-22T07:21:05.717+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04773 | 2022-09-22T07:18:45.727+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04983 | 2022-09-22T07:18:45.677+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04069 | 2022-09-22T07:18:31.837+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05029 | 2022-09-22T07:18:23.957+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04984 | 2022-09-22T07:15:48.15+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04720 | 2022-09-22T07:15:44.87+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03875 | 2022-09-22T07:15:23.863+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04702 | 2022-09-22T07:15:13.127+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03926 | 2022-09-22T07:00:50.103+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04751 | 2022-09-22T07:00:25.03+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03980 | 2022-09-22T07:00:19.32+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04469 | 2022-09-22T06:59:02.577+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04653 | 2022-09-22T06:58:42.43+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05185 | 2022-09-22T06:58:14.467+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03746 | 2022-09-22T06:58:04.347+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05196 | 2022-09-22T06:56:49.977+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04526 | 2022-09-22T06:56:33.33+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04736 | 2022-09-22T06:56:03.157+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04889 | 2022-09-22T06:55:46.353+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05259 | 2022-09-22T06:54:37.107+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04528 | 2022-09-22T06:54:21.55+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05228 | 2022-09-22T06:53:35.987+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05284 | 2022-09-22T06:53:25.68+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05018 | 2022-09-22T06:52:15.353+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05096 | 2022-09-22T06:51:58.797+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04254 | 2022-09-22T06:51:23.21+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05228 | 2022-09-22T06:51:15.033+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04186 | 2022-09-22T06:50:07.917+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04695 | 2022-09-22T06:49:45.91+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04842 | 2022-09-22T06:49:29.483+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04571 | 2022-09-22T06:48:56.577+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04726 | 2022-09-22T06:47:48.36+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04100 | 2022-09-22T06:47:21.187+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04158 | 2022-09-22T06:46:59.267+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04837 | 2022-09-22T06:46:34.947+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04739 | 2022-09-22T06:45:12.897+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04142 | 2022-09-22T06:44:57.497+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04106 | 2022-09-22T06:44:45.697+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03872 | 2022-09-22T06:44:21.26+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03853 | 2022-09-22T06:42:00.94+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04564 | 2022-09-22T06:41:41.25+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03940 | 2022-09-22T06:41:31.393+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04601 | 2022-09-22T06:40:24.48+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04365 | 2022-09-22T06:38:23.41+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03915 | 2022-09-22T06:38:21.11+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03773 | 2022-09-22T06:37:59.68+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04358 | 2022-09-22T06:37:41.027+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04776 | 2022-09-22T06:33:42.543+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04840 | 2022-09-22T06:33:40.393+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04323 | 2022-09-22T06:33:27.573+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04274 | 2022-09-22T06:33:18.67+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03938 | 2022-09-22T06:29:24.017+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04465 | 2022-09-22T06:29:23.357+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05128 | 2022-09-22T06:29:07.677+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04489 | 2022-09-22T06:29:00.887+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03631 | 2022-09-22T06:25:54.19+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03862 | 2022-09-22T06:25:52.68+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04383 | 2022-09-22T06:25:29.4+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04952 | 2022-09-22T06:25:20.42+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04872 | 2022-09-22T06:23:01.523+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05247 | 2022-09-22T06:22:58.267+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03808 | 2022-09-22T06:22:33.177+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04777 | 2022-09-22T06:22:12.74+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04708 | 2022-09-22T06:20:19.567+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03833 | 2022-09-22T06:20:18.937+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04754 | 2022-09-22T06:20:05.187+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04229 | 2022-09-22T06:19:55.847+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04295 | 2022-09-22T06:17:43.7+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03971 | 2022-09-22T06:17:37.86+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04245 | 2022-09-22T06:17:12.407+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04543 | 2022-09-22T06:17:02.657+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04107 | 2022-09-22T06:15:17.353+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03923 | 2022-09-22T06:15:15.293+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04987 | 2022-09-22T06:14:58.38+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03946 | 2022-09-22T06:14:46.177+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04239 | 2022-09-22T06:12:48.72+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03968 | 2022-09-22T06:12:48.56+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03763 | 2022-09-22T06:12:28.09+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04376 | 2022-09-22T06:12:19.39+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04429 | 2022-09-22T06:09:55.827+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04553 | 2022-09-22T06:09:55.507+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04006 | 2022-09-22T06:09:36.807+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04675 | 2022-09-22T06:09:29.097+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04879 | 2022-09-22T06:06:54.607+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04765 | 2022-09-22T06:06:50.52+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05194 | 2022-09-22T06:06:34.91+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04608 | 2022-09-22T06:06:25.42+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04835 | 2022-09-22T06:03:41.913+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04782 | 2022-09-22T06:03:41.483+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04619 | 2022-09-22T06:03:13.057+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04664 | 2022-09-22T06:03:02.787+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03982 | 2022-09-22T06:01:30.567+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05146 | 2022-09-22T06:01:25.39+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05296 | 2022-09-22T06:00:55.947+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04662 | 2022-09-22T06:00:47.887+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04095 | 2022-09-22T05:59:11.343+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05068 | 2022-09-22T05:58:57.01+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04375 | 2022-09-22T05:58:39.833+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05105 | 2022-09-22T05:58:31.347+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03649 | 2022-09-22T05:56:14.333+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04585 | 2022-09-22T05:56:12.183+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04632 | 2022-09-22T05:55:58.393+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03271 | 2022-09-22T05:55:47.74+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04010 | 2022-09-22T05:52:47.88+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05086 | 2022-09-22T05:52:45.14+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04730 | 2022-09-22T05:51:44.317+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04624 | 2022-09-22T05:51:32.71+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04042 | 2022-09-22T05:49:49.543+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04529 | 2022-09-22T05:49:45.713+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04540 | 2022-09-22T05:49:15.417+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04834 | 2022-09-22T05:49:08.11+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04513 | 2022-09-22T05:46:17.587+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04706 | 2022-09-22T05:46:05.1+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04137 | 2022-09-22T05:45:43.237+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04617 | 2022-09-22T05:45:29.833+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04605 | 2022-09-22T05:42:47.707+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04730 | 2022-09-22T05:42:12.017+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04298 | 2022-09-22T05:42:08.94+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03943 | 2022-09-22T05:41:57.8+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04594 | 2022-09-22T05:40:07.337+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04738 | 2022-09-22T05:39:56.177+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03689 | 2022-09-22T05:39:35.447+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04101 | 2022-09-22T05:39:27.807+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04719 | 2022-09-22T05:34:54.013+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03991 | 2022-09-22T05:34:49.073+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05040 | 2022-09-22T05:34:29.937+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04728 | 2022-09-22T05:34:16.887+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04242 | 2022-09-22T05:31:30.5+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04020 | 2022-09-22T05:31:25.28+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04697 | 2022-09-22T05:31:00.98+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04251 | 2022-09-22T05:30:49.673+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03180 | 2022-09-22T05:28:13.267+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04622 | 2022-09-22T05:27:58.6+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04771 | 2022-09-22T05:27:34.117+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04897 | 2022-09-22T05:27:18.477+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04233 | 2022-09-22T05:24:38.78+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04169 | 2022-09-22T05:24:34.63+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04048 | 2022-09-22T05:24:04.62+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04749 | 2022-09-22T05:23:55.76+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04001 | 2022-09-22T05:22:03.167+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04486 | 2022-09-22T05:21:55.957+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03772 | 2022-09-22T05:21:38.39+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03981 | 2022-09-22T05:21:13.567+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04384 | 2022-09-22T05:17:54.75+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03868 | 2022-09-22T05:17:52.47+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04031 | 2022-09-22T05:17:36.4+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03972 | 2022-09-22T05:17:26.79+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04330 | 2022-09-22T05:13:44.647+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04034 | 2022-09-22T05:13:41.157+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04770 | 2022-09-22T05:13:25.767+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04641 | 2022-09-22T05:13:09.713+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04898 | 2022-09-22T04:28:57.01+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03924 | 2022-09-22T04:28:52.357+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04531 | 2022-09-22T04:28:47.767+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04745 | 2022-09-22T04:28:45.297+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04159 | 2022-09-22T03:00:56.78+00:00 | FALSE | 255 | 27 | . | 704
16045256 | 17858 | 1 | 2905E03951 | 2022-09-22T02:59:23.147+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04731 | 2022-09-22T02:58:39.517+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04200 | 2022-09-22T02:58:22.187+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04159 | 2022-09-22T02:57:15.447+00:00 | FALSE | 3 | 20 | . | 704
16045256 | 17858 | 1 | 2905E04286 | 2022-09-22T02:56:55.563+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04484 | 2022-09-22T02:54:58.147+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04682 | 2022-09-22T02:54:29.673+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04032 | 2022-09-22T02:54:17.83+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04688 | 2022-09-22T02:53:37.663+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04700 | 2022-09-22T02:52:36.067+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05135 | 2022-09-22T02:51:25.833+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04948 | 2022-09-22T02:50:58.38+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04159 | 2022-09-22T02:50:45.16+00:00 | FALSE | 2 | 9 | . | 704
16045256 | 17858 | 3 | 2905E04458 | 2022-09-22T02:50:05.457+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04535 | 2022-09-22T02:49:19.54+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04286 | 2022-09-22T02:47:41.877+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04159 | 2022-09-22T02:47:33.507+00:00 | FALSE | 4 | 9 | . | 704
16045256 | 17858 | 3 | 2905E03705 | 2022-09-22T02:47:28.75+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04589 | 2022-09-22T02:46:53.967+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04380 | 2022-09-22T02:46:27.29+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04159 | 2022-09-22T02:44:54.233+00:00 | FALSE | 255 | 9 | . | 704
16045256 | 17858 | 3 | 2905E04348 | 2022-09-22T02:44:50.607+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03925 | 2022-09-22T02:44:25.163+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03903 | 2022-09-22T02:43:57.597+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04635 | 2022-09-22T02:43:38.96+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04159 | 2022-09-22T02:42:28.057+00:00 | FALSE | 3 | 20 | . | 704
16045256 | 17858 | 1 | 2905E03969 | 2022-09-22T02:41:41.767+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04348 | 2022-09-22T02:41:17.877+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04846 | 2022-09-22T02:41:12.047+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04028 | 2022-09-22T02:40:37.46+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04159 | 2022-09-22T02:39:30.31+00:00 | FALSE | 2 | 20 | . | 704
16045256 | 17858 | 2 | 2905E04045 | 2022-09-22T02:29:37.933+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E05320 | 2022-09-22T02:28:36.567+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04828 | 2022-09-22T02:28:06.713+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05110 | 2022-09-22T02:27:44.78+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04559 | 2022-09-22T02:27:09.337+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04045 | 2022-09-22T02:25:43.503+00:00 | FALSE | 2 | 30 | . | 704
16045256 | 17858 | 4 | 2905E05126 | 2022-09-22T02:25:41.317+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04717 | 2022-09-22T02:25:24.08+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03895 | 2022-09-22T02:24:53.35+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05013 | 2022-09-22T02:23:56.367+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05138 | 2022-09-22T02:23:17.147+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05325 | 2022-09-22T02:23:00.647+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04982 | 2022-09-22T02:22:26.17+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05108 | 2022-09-22T02:21:40.243+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04207 | 2022-09-22T02:21:06.193+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03974 | 2022-09-22T02:20:39.107+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04119 | 2022-09-22T02:19:24.313+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04199 | 2022-09-22T02:19:21.723+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05317 | 2022-09-22T02:18:18.08+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03852 | 2022-09-22T02:17:47.127+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04631 | 2022-09-22T02:17:31.617+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04584 | 2022-09-22T02:16:54.437+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04714 | 2022-09-22T02:14:51.327+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04381 | 2022-09-22T02:14:40.667+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04953 | 2022-09-22T02:14:13.437+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04045 | 2022-09-22T02:13:06.157+00:00 | FALSE | 2 | 30 | . | 704
16045256 | 17858 | 2 | 2905E05225 | 2022-09-22T01:34:27.68+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E05007 | 2022-09-22T01:33:59.11+00:00 | FALSE | 3 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04957 | 2022-09-22T01:30:41.78+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05007 | 2022-09-22T01:30:38.143+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E05225 | 2022-09-22T01:30:01.74+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 2 | 2905E04957 | 2022-09-22T01:28:18.053+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E05007 | 2022-09-22T01:27:42.947+00:00 | FALSE | 2 | 30 | . | 704
16045256 | 17858 | 4 | 2905E05053 | 2022-09-22T01:27:15.367+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 3 | 2905E05225 | 2022-09-22T01:26:52.357+00:00 | FALSE | 3 | 30 | . | 704
16045256 | 17858 | 2 | 2905E05124 | 2022-09-22T01:26:20.57+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04705 | 2022-09-22T01:25:45.207+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04139 | 2022-09-22T01:25:18.387+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03507 | 2022-09-22T01:24:56.327+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03586 | 2022-09-22T01:23:42.323+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04431 | 2022-09-22T01:23:08.82+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03438 | 2022-09-22T01:22:42.717+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05155 | 2022-09-22T01:22:17.937+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04139 | 2022-09-22T01:21:11.683+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E05053 | 2022-09-22T01:20:33.643+00:00 | FALSE | 2 | 30 | . | 704
16045256 | 17858 | 4 | 2905E03135 | 2022-09-22T01:19:58.783+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03714 | 2022-09-22T01:19:35.723+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05189 | 2022-09-22T01:19:13.657+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03665 | 2022-09-22T01:18:42.173+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05053 | 2022-09-22T01:17:33.657+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04814 | 2022-09-22T01:17:18.823+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05017 | 2022-09-22T01:16:51.997+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05285 | 2022-09-22T01:16:23.2+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04932 | 2022-09-22T01:15:38.627+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04732 | 2022-09-22T01:15:02.833+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04814 | 2022-09-22T01:14:14.61+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E05336 | 2022-09-22T01:13:37.823+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04916 | 2022-09-22T01:13:09.153+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04896 | 2022-09-22T01:12:30.273+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04732 | 2022-09-22T01:12:25.467+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 2 | 2905E04904 | 2022-09-22T01:09:06.98+00:00 | FALSE | 255 | 15 | . | 704
16045256 | 17858 | 2 | 2905E03550 | 2022-09-22T01:03:30.047+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04904 | 2022-09-22T01:02:15.7+00:00 | FALSE | 3 | 15 | . | 704
16045256 | 17858 | 1 | 2905E03698 | 2022-09-22T01:01:49.763+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04221 | 2022-09-22T01:01:08.757+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04856 | 2022-09-22T01:00:59.877+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03818 | 2022-09-22T00:59:11.617+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04904 | 2022-09-22T00:59:00.713+00:00 | FALSE | 2 | 15 | . | 704
16045256 | 17858 | 4 | 2905E04146 | 2022-09-22T00:58:47.267+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04734 | 2022-09-22T00:58:11.947+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03896 | 2022-09-22T00:56:59.38+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03343 | 2022-09-22T00:56:01.097+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05078 | 2022-09-22T00:55:45.107+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03693 | 2022-09-22T00:55:03.47+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04878 | 2022-09-22T00:54:38.343+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04989 | 2022-09-22T00:53:39.377+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03784 | 2022-09-22T00:53:15.957+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04832 | 2022-09-22T00:52:40.69+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04534 | 2022-09-22T00:52:11.78+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04929 | 2022-09-22T00:51:28.937+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05023 | 2022-09-22T00:50:42.307+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04947 | 2022-09-22T00:50:15.247+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05127 | 2022-09-22T00:49:45.89+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05074 | 2022-09-22T00:49:11.47+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E02667 | 2022-09-22T00:48:03.41+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05011 | 2022-09-22T00:47:32.177+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05084 | 2022-09-22T00:46:49.67+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04908 | 2022-09-22T00:46:19.45+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05097 | 2022-09-22T00:45:38.53+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03385 | 2022-09-22T00:44:58.093+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04679 | 2022-09-22T00:44:33.367+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05217 | 2022-09-22T00:43:56.917+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04667 | 2022-09-22T00:43:18.11+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04378 | 2022-09-22T00:42:43.697+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04848 | 2022-09-22T00:42:12.937+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04679 | 2022-09-22T00:41:31.913+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04679 | 2022-09-22T00:40:27.893+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E03691 | 2022-09-22T00:40:14.79+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04690 | 2022-09-22T00:39:52.263+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03643 | 2022-09-22T00:39:15.863+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04848 | 2022-09-22T00:38:12.59+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04669 | 2022-09-22T00:37:59.497+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03733 | 2022-09-22T00:37:37.713+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04250 | 2022-09-22T00:36:59.217+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E05051 | 2022-09-22T00:36:17.963+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04748 | 2022-09-22T00:35:38.597+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04939 | 2022-09-22T00:35:15.983+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03847 | 2022-09-22T00:34:44.097+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04322 | 2022-09-22T00:33:52.917+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04615 | 2022-09-22T00:32:57.35+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04906 | 2022-09-22T00:32:33.34+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04946 | 2022-09-22T00:32:05.757+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04134 | 2022-09-22T00:31:27.513+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E05022 | 2022-09-22T00:30:25.38+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04671 | 2022-09-22T00:30:02.903+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04742 | 2022-09-22T00:29:27.967+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03134 | 2022-09-22T00:28:27.22+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04744 | 2022-09-22T00:27:55.623+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04860 | 2022-09-22T00:27:23.43+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05028 | 2022-09-22T00:26:53.167+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04822 | 2022-09-22T00:25:58.873+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04985 | 2022-09-22T00:25:27.997+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04372 | 2022-09-22T00:24:59.97+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05316 | 2022-09-22T00:24:38.063+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04019 | 2022-09-22T00:23:30.137+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04787 | 2022-09-22T00:23:00.017+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05243 | 2022-09-22T00:22:31.09+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03661 | 2022-09-22T00:22:22.543+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04950 | 2022-09-22T00:21:00.837+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04054 | 2022-09-22T00:20:34.7+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03555 | 2022-09-22T00:20:06.333+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04054 | 2022-09-22T00:18:12.337+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04960 | 2022-09-22T00:18:09.467+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04723 | 2022-09-22T00:17:34.517+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04689 | 2022-09-22T00:15:14.793+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04072 | 2022-09-22T00:15:04.94+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04265 | 2022-09-22T00:15:02.567+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04960 | 2022-09-22T00:13:36.753+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 2 | 2905E04723 | 2022-09-22T00:12:05.347+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E03979 | 2022-09-21T11:11:20.177+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04182 | 2022-09-21T11:11:10.397+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03874 | 2022-09-21T11:09:31.563+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04661 | 2022-09-21T11:09:13.833+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04692 | 2022-09-21T11:09:05.223+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03272 | 2022-09-21T11:08:57.307+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03846 | 2022-09-21T11:07:15.987+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03473 | 2022-09-21T11:07:01.123+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05129 | 2022-09-21T11:06:50.963+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04674 | 2022-09-21T11:06:45.197+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04737 | 2022-09-21T11:04:58.233+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04370 | 2022-09-21T11:04:44.087+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05329 | 2022-09-21T11:04:32.027+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04572 | 2022-09-21T11:04:27.877+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04560 | 2022-09-21T11:02:27.877+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04673 | 2022-09-21T11:02:18.717+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04942 | 2022-09-21T11:02:13.983+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03541 | 2022-09-21T11:00:04.59+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04936 | 2022-09-21T10:59:56.81+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03472 | 2022-09-21T10:59:49.357+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03886 | 2022-09-21T10:57:44.787+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03143 | 2022-09-21T10:57:37.837+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04243 | 2022-09-21T10:57:32.533+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04958 | 2022-09-21T10:54:13.67+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E05136 | 2022-09-21T10:54:03.537+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04805 | 2022-09-21T10:53:58.347+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05292 | 2022-09-21T10:51:49.86+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04678 | 2022-09-21T10:51:43.08+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04847 | 2022-09-21T10:51:34.47+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05067 | 2022-09-21T10:47:39.477+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04709 | 2022-09-21T10:47:28.197+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04833 | 2022-09-21T10:47:21.07+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04839 | 2022-09-21T10:45:29.487+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04850 | 2022-09-21T10:45:21.94+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04647 | 2022-09-21T10:45:11.25+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E05197 | 2022-09-21T10:43:16.473+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04646 | 2022-09-21T10:43:03.403+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04949 | 2022-09-21T10:42:59.203+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04647 | 2022-09-21T10:42:46.273+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E05027 | 2022-09-21T10:34:43.093+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03962 | 2022-09-21T10:34:32.527+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04600 | 2022-09-21T10:33:42.27+00:00 | FALSE | 4 | 15 | . | 704
16045256 | 17858 | 3 | 2905E05180 | 2022-09-21T10:32:24.35+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04786 | 2022-09-21T10:31:26.557+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04600 | 2022-09-21T10:26:00.293+00:00 | FALSE | 2 | 15 | . | 704
16045256 | 17858 | 4 | 2905E04546 | 2022-09-21T10:25:11.75+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04786 | 2022-09-21T10:24:40.12+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04685 | 2022-09-21T10:24:30.053+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04894 | 2022-09-21T10:22:54.33+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04456 | 2022-09-21T10:22:27.403+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04410 | 2022-09-21T10:21:57.67+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04446 | 2022-09-21T10:21:27.233+00:00 | FALSE | 2 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04361 | 2022-09-21T10:20:03.973+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04336 | 2022-09-21T10:20:02.01+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04340 | 2022-09-21T10:19:48.11+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04487 | 2022-09-21T10:19:38.143+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04446 | 2022-09-21T10:17:05.607+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04327 | 2022-09-21T10:17:04.697+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04566 | 2022-09-21T10:16:56.777+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04487 | 2022-09-21T10:16:53.997+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04481 | 2022-09-21T10:14:56.26+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04566 | 2022-09-21T10:14:25.18+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04545 | 2022-09-21T10:14:19.44+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04354 | 2022-09-21T10:14:10.143+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04480 | 2022-09-21T10:12:48.263+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04567 | 2022-09-21T10:12:41.743+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04593 | 2022-09-21T10:12:07.89+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04354 | 2022-09-21T10:11:55.677+00:00 | FALSE | 3 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04144 | 2022-09-21T10:10:20.19+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04762 | 2022-09-21T10:10:02.283+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04575 | 2022-09-21T10:09:55.283+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04593 | 2022-09-21T10:09:49.163+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E03999 | 2022-09-21T10:07:42.003+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04520 | 2022-09-21T10:07:37.203+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04459 | 2022-09-21T10:07:22.973+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04349 | 2022-09-21T10:07:16.427+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04104 | 2022-09-21T10:05:36.183+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03976 | 2022-09-21T10:05:12.787+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04457 | 2022-09-21T10:05:06.06+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04349 | 2022-09-21T10:05:01.823+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 2 | 2905E04240 | 2022-09-21T10:03:00.277+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03851 | 2022-09-21T10:02:59.91+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04507 | 2022-09-21T10:02:46.907+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04547 | 2022-09-21T10:02:39.647+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03883 | 2022-09-21T09:58:55.317+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E01114 | 2022-09-21T09:58:24.23+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03963 | 2022-09-21T09:58:06.273+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03777 | 2022-09-21T09:57:56.823+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04021 | 2022-09-21T09:56:33.62+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E01267 | 2022-09-21T09:55:50.077+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03777 | 2022-09-21T09:55:35.987+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04000 | 2022-09-21T09:55:34.657+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04035 | 2022-09-21T09:53:16.803+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04050 | 2022-09-21T09:53:14.523+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04012 | 2022-09-21T09:52:51.357+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04521 | 2022-09-21T09:52:41.667+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E01228 | 2022-09-21T09:51:04.36+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E00942 | 2022-09-21T09:50:56.62+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E01152 | 2022-09-21T09:50:20.437+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E01265 | 2022-09-21T09:50:14.02+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04407 | 2022-09-21T09:48:37.383+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E01440 | 2022-09-21T09:48:36.383+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04576 | 2022-09-21T09:48:05.383+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E01048 | 2022-09-21T09:48:00.503+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01142 | 2022-09-21T09:46:19.437+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04660 | 2022-09-21T09:46:15.7+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E00953 | 2022-09-21T09:45:52.207+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04241 | 2022-09-21T09:45:45.187+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04391 | 2022-09-21T09:43:26.287+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04030 | 2022-09-21T09:43:24.457+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 1 | 2905E03931 | 2022-09-21T09:43:20.593+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04241 | 2022-09-21T09:43:11.953+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04321 | 2022-09-21T09:41:36.687+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04488 | 2022-09-21T09:41:32.177+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04224 | 2022-09-21T09:41:16.213+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04477 | 2022-09-21T09:41:00.55+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04147 | 2022-09-21T09:39:12.23+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04577 | 2022-09-21T09:38:53.293+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04321 | 2022-09-21T09:38:38.163+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04030 | 2022-09-21T09:38:18.46+00:00 | FALSE | 2 | 30 | . | 704
16045256 | 17858 | 4 | 2905E03934 | 2022-09-21T09:36:49.353+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04583 | 2022-09-21T09:36:33.833+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04002 | 2022-09-21T09:36:17.793+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04030 | 2022-09-21T09:36:17.023+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E03882 | 2022-09-21T09:33:30.8+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E03881 | 2022-09-21T09:33:29.67+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04788 | 2022-09-21T09:33:14.36+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04437 | 2022-09-21T09:33:07.753+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04132 | 2022-09-21T08:55:33.143+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04956 | 2022-09-21T08:55:24.123+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04655 | 2022-09-21T08:54:55.063+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04437 | 2022-09-21T08:54:37.1+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E05119 | 2022-09-21T08:53:22.113+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03873 | 2022-09-21T08:53:15.69+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03739 | 2022-09-21T08:52:47.753+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04451 | 2022-09-21T08:52:47.533+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04628 | 2022-09-21T08:51:08.47+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E03913 | 2022-09-21T08:50:58.55+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04831 | 2022-09-21T08:50:45.783+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04382 | 2022-09-21T08:50:09.437+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04639 | 2022-09-21T08:49:03.91+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04820 | 2022-09-21T08:48:54.827+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04036 | 2022-09-21T08:48:34.36+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04637 | 2022-09-21T08:46:18.787+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04093 | 2022-09-21T08:46:07.407+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04320 | 2022-09-21T08:45:56.623+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04036 | 2022-09-21T08:45:18.153+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04992 | 2022-09-21T08:44:06.277+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 1 | 2905E03811 | 2022-09-21T08:43:36.413+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04831 | 2022-09-21T08:43:32.573+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04093 | 2022-09-21T08:43:21.787+00:00 | FALSE | 3 | 30 | . | 704
16045256 | 17858 | 2 | 2905E04373 | 2022-09-21T08:39:10.81+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04654 | 2022-09-21T08:38:57.74+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04597 | 2022-09-21T08:38:54.79+00:00 | FALSE | 4 | 20 | . | 704
16045256 | 17858 | 1 | 2905E04310 | 2022-09-21T08:38:27.067+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04164 | 2022-09-21T08:36:55.817+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 4 | 2905E03884 | 2022-09-21T08:36:52.417+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04448 | 2022-09-21T08:36:31.007+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04780 | 2022-09-21T08:36:18.217+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04027 | 2022-09-21T08:33:28.35+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04597 | 2022-09-21T08:33:17.46+00:00 | FALSE | 2 | 59 | . | 704
16045256 | 17858 | 3 | 2905E03820 | 2022-09-21T08:33:10.08+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04310 | 2022-09-21T08:32:56.097+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 4 | 2905E04710 | 2022-09-21T08:30:33.323+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04504 | 2022-09-21T08:30:30.303+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E04557 | 2022-09-21T08:30:14.803+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E04524 | 2022-09-21T08:29:59.913+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E04557 | 2022-09-21T08:27:58.627+00:00 | FALSE | 4 | 30 | . | 704
16045256 | 17858 | 1 | 2905E04445 | 2022-09-21T08:27:49.147+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 2 | 2905E04524 | 2022-09-21T08:27:42.197+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E04597 | 2022-09-21T08:27:34.637+00:00 | FALSE | 3 | 20 | . | 704
16045256 | 17858 | 1 | 2905E01005 | 2022-09-21T08:19:02.547+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E00972 | 2022-09-21T08:18:53.047+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01259 | 2022-09-21T08:17:33.48+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E00996 | 2022-09-21T08:17:15.813+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E01100 | 2022-09-21T08:17:00.083+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E01147 | 2022-09-21T08:16:49.933+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E00939 | 2022-09-21T08:15:31.287+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E01100 | 2022-09-21T08:15:00.353+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 3 | 2905E01725 | 2022-09-21T08:14:42.247+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E01013 | 2022-09-21T08:14:12.367+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01253 | 2022-09-21T08:13:20.357+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E01101 | 2022-09-21T08:13:17.597+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E01085 | 2022-09-21T08:12:37.747+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E01291 | 2022-09-21T08:12:07.743+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01289 | 2022-09-21T08:11:04.487+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E01127 | 2022-09-21T08:11:03.037+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E00975 | 2022-09-21T08:10:33.29+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E00948 | 2022-09-21T08:10:04.22+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01014 | 2022-09-21T08:08:53.293+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E01153 | 2022-09-21T08:08:51.003+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E01069 | 2022-09-21T08:08:18.067+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 1 | 2905E00624 | 2022-09-21T08:07:54.587+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01363 | 2022-09-21T08:06:16.933+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E00624 | 2022-09-21T08:05:36.613+00:00 | FALSE | 255 | 30 | . | 704
16045256 | 17858 | 1 | 2905E00931 | 2022-09-21T08:04:24.747+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 3 | 2905E01273 | 2022-09-21T08:04:18.707+00:00 | TRUE | 3 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01216 | 2022-09-21T08:04:02.097+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E00861 | 2022-09-21T08:03:58.017+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 1 | 2905E00925 | 2022-09-21T08:02:20.873+00:00 | TRUE | 2 | 0 | . | 704
16045256 | 17858 | 4 | 2905E01354 | 2022-09-21T08:01:37.177+00:00 | TRUE | 4 | 0 | . | 704
16045256 | 17858 | 2 | 2905E00711 | 2022-09-21T08:01:36.517+00:00 | TRUE | 255 | 0 | . | 704
16045256 | 17858 | 3 | 2905E00861 | 2022-09-21T08:01:07.36+00:00 | FALSE | 3 | 59 | . | 704
16045256 | 17858 | 1 | 2905E01423 | 2022-09-21T08:00:11.117+00:00 | TRUE | 2 | 0 | . | 704

## Retest Item Detail
JobNo | ItemName Type | Process Type | Test Type | Serial Number | Failitem | Failitem name | Fail Value | TDate Time (UTC) | STATION | FixtureID2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 71 | D104 PD dark noise OHR dev | 208.2552 | 2022-09-22T06:33:23.367+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 71 | D104 PD dark noise OHR dev | 204.5979 | 2022-09-23T03:37:20.387+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 71 | D104 PD dark noise OHR dev | 214.5768 | 2022-09-23T03:40:59.313+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 80 | D110 – 940nm Brightness (5mA) Measure diff with V_ref | -0.739 | 2022-09-22T06:40:40.2+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 69 | D102 PD dark noise OHR dev | 209.3026 | 2022-09-22T10:55:15.22+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 69 | D102 PD dark noise OHR dev | 221.5885 | 2022-09-23T00:23:21.417+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 41 | Barometer Temperature | -270.8047 | 2022-09-23T00:19:42.68+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04594 | 41 | Barometer Temperature | -999 | 2022-09-23T01:04:20.657+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05247 | 41 | Barometer Temperature | -999 | 2022-09-22T08:21:18.947+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05247 | 41 | Barometer Temperature | -999 | 2022-09-22T08:22:48.5+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05247 | 41 | Barometer Temperature | 22.39063 | 2022-09-23T00:18:06.51+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03940 | 98 | D109 – 540nm Brightness (20mA) | -999 | 2022-09-22T07:56:29.03+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03940 | 41 | Barometer Temperature | -999 | 2022-09-22T07:59:15.9+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03940 | 41 | Barometer Temperature | 22.1875 | 2022-09-23T00:14:24.587+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04782 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:26:59.37+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04782 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:35:11.287+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04782 | 88 | AFE Power Consumption data | -999 | 2022-09-22T11:03:10.417+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04782 | 88 | AFE Power Consumption data | 2517.43 | 2022-09-23T00:35:36.377+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03991 | 41 | Barometer Temperature | -999 | 2022-09-22T06:34:07.393+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03991 | 41 | Barometer Temperature | -999 | 2022-09-22T06:42:35.157+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03991 | 41 | Barometer Temperature | 24.22656 | 2022-09-22T10:55:21.883+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04751 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:16:56.317+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04751 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:19:51.953+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04751 | 88 | AFE Power Consumption data | 3149.75 | 2022-09-22T10:47:50.217+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04811 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:13:12.82+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04811 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:39:40.967+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04811 | 88 | AFE Power Consumption data | -999 | 2022-09-22T11:11:47.013+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04811 | 88 | AFE Power Consumption data | 2947.172 | 2022-09-23T00:29:39.417+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04436 | 49 | AFE Part ID Read | -999 | 2022-09-22T09:50:17.367+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04436 | 49 | AFE Part ID Read | 808800 | 2022-09-22T09:58:17.497+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03942 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T08:48:08.83+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03942 | 46 | Temperature Sensor Temperature | 25.19 | 2022-09-22T09:40:55.357+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04543 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:24:52.833+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04543 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:28:32.693+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04543 | 88 | AFE Power Consumption data | 2461.743 | 2022-09-23T00:30:32.993+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03773 | 41 | Barometer Temperature | -999 | 2022-09-22T07:59:37.237+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03773 | 41 | Barometer Temperature | -999 | 2022-09-22T08:00:50.47+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03773 | 41 | Barometer Temperature | 22.46875 | 2022-09-23T00:17:39.287+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04837 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:50:04.603+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04837 | 88 | AFE Power Consumption data | 3091.154 | 2022-09-22T07:54:15.333+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04749 | 73 | D109 – 630nm Brightness (20mA) Measure diff with V_ref | -0.548 | 2022-09-22T05:40:27.43+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04749 | 73 | D109 – 630nm Brightness (20mA) Measure diff with V_ref | 830.149 | 2022-09-22T06:11:04.127+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04575 | 41 | Barometer Temperature | -999 | 2022-09-22T01:18:22.677+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04575 | 41 | Barometer Temperature | 24.77344 | 2022-09-22T04:04:24.493+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04906 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:56:09.883+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04906 | 88 | AFE Power Consumption data | 3033.059 | 2022-09-22T04:01:06.527+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03811 | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | -2.088 | 2022-09-22T03:09:38.957+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03811 | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | 796.109 | 2022-09-22T05:19:56.117+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03820 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:05:15.9+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03820 | 88 | AFE Power Consumption data | 3193.562 | 2022-09-22T05:18:43.603+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04373 | 55 | D106 – 540nm Brightness (20mA) Measure diff with V_ref | 0.055 | 2022-09-22T02:59:38.127+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04373 | 55 | D106 – 540nm Brightness (20mA) Measure diff with V_ref | 887.904 | 2022-09-22T03:02:12.947+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03134 | 41 | Barometer Temperature | -999 | 2022-09-22T02:25:29.783+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03134 | 41 | Barometer Temperature | 24.16406 | 2022-09-22T02:34:22.08+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04451 | 41 | Barometer Temperature | -999 | 2022-09-22T02:10:52.267+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04451 | 41 | Barometer Temperature | -999 | 2022-09-22T02:18:18.743+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04451 | 41 | Barometer Temperature | -999 | 2022-09-22T06:43:19.22+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04451 | 41 | Barometer Temperature | -999 | 2022-09-23T00:42:10.31+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04451 | 41 | Barometer Temperature | -999 | 2022-09-23T00:43:43.477+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04451 | 41 | Barometer Temperature | 23.72656 | 2022-09-23T03:36:34.487+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04690 | 88 | AFE Power Consumption data | -999 | 2022-09-22T04:11:08.807+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04690 | 88 | AFE Power Consumption data | -999 | 2022-09-23T00:46:46.197+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04690 | 41 | Barometer Temperature | -999 | 2022-09-22T06:45:43.173+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04690 | 41 | Barometer Temperature | -999 | 2022-09-23T00:48:55.513+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04690 | 41 | Barometer Temperature | 23.79688 | 2022-09-23T03:39:40.013+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03952 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:02:12.647+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03952 | 88 | AFE Power Consumption data | 3150.57 | 2022-09-22T10:38:24.117+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04578 | 3 | Fixture ID | -999 | 2022-09-22T09:44:29.797+00:00 | T1_5PR1AT05
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04578 | 3 | Fixture ID | 1 | 2022-09-22T09:54:41.403+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03763 | 41 | Barometer Temperature | -999 | 2022-09-22T08:34:48.223+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03763 | 41 | Barometer Temperature | -999 | 2022-09-22T08:39:50.517+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03763 | 41 | Barometer Temperature | 24.80469 | 2022-09-22T10:59:46.343+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04952 | 74 | D109 – 660nm Brightness (20mA) Measure diff with V_ref | 2.772 | 2022-09-22T08:16:32.963+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04952 | 74 | D109 – 660nm Brightness (20mA) Measure diff with V_ref | 1011.671 | 2022-09-22T08:20:05.42+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04842 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:44:57.28+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04842 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:46:20.617+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04842 | 88 | AFE Power Consumption data | -999 | 2022-09-23T00:25:51.813+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04842 | 88 | AFE Power Consumption data | -999 | 2022-09-23T00:40:21.873+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04011 | 41 | Barometer Temperature | -999 | 2022-09-22T07:34:02.027+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04011 | 41 | Barometer Temperature | 24.5 | 2022-09-22T07:40:00.2+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03851 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:24:55.503+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03851 | 34 | AFE Power Consumption Test | 0.196 | 2022-09-22T04:03:47.087+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03851 | 34 | AFE Power Consumption Test | 0.152 | 2022-09-22T06:46:20.07+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04958 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:41:33.883+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04958 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:46:08.233+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04958 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:36:23.48+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04958 | 41 | Barometer Temperature | -999 | 2022-09-22T01:40:45.493+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04958 | 41 | Barometer Temperature | -999 | 2022-09-23T00:41:45.423+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04958 | 41 | Barometer Temperature | 23.21875 | 2022-09-23T00:45:16.107+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04958 | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | -3.223 | 2022-09-22T06:44:28.483+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03924 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:50:40.773+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03924 | 41 | Barometer Temperature | -999 | 2022-09-22T05:52:13.133+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03924 | 41 | Barometer Temperature | 24.28906 | 2022-09-22T06:27:47.23+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04298 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:13:29.75+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04298 | 88 | AFE Power Consumption data | 2979.436 | 2022-09-22T06:16:21.213+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04947 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:43:50.707+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04947 | 88 | AFE Power Consumption data | 2937.749 | 2022-09-22T05:22:46.153+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04832 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:44:08.62+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04832 | 88 | AFE Power Consumption data | 3328.929 | 2022-09-22T05:33:25.247+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04072 | 88 | AFE Power Consumption data | -999 | 2022-09-22T02:12:21.733+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04072 | 88 | AFE Power Consumption data | 2411.827 | 2022-09-22T02:16:19.713+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01152 | 41 | Barometer Temperature | -999 | 2022-09-22T01:39:00.773+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01152 | 41 | Barometer Temperature | 24.59375 | 2022-09-22T03:04:10.617+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04737 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:29:58.153+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04737 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:49:40.03+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04737 | 88 | AFE Power Consumption data | 2410.291 | 2022-09-22T01:36:34.397+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01253 | 4 | USB Connection 1 |  | 2022-09-21T09:59:20.477+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01253 | 88 | AFE Power Consumption data | -999 | 2022-09-21T10:03:15.397+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01253 | 88 | AFE Power Consumption data | 3064.912 | 2022-09-21T10:06:40.867+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04139 | 10 | D103 Response avg | 16.64274 | 2022-09-22T03:41:32.623+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04139 | 10 | D103 Response avg | 16.1421 | 2022-09-22T05:05:36.57+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04139 | 10 | D103 Response avg | 16.68638 | 2022-09-22T06:55:16.697+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04139 | 10 | D103 Response avg | 16.65129 | 2022-09-22T06:58:42.833+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04139 | 10 | D103 Response avg | 17.03222 | 2022-09-22T10:52:25.797+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04139 | 10 | D103 Response avg | 15.71942 | 2022-09-23T00:23:53.827+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03926 | 41 | Barometer Temperature | -999 | 2022-09-22T07:15:01.797+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03926 | 41 | Barometer Temperature | -999 | 2022-09-22T07:18:07.33+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03926 | 41 | Barometer Temperature | 24.77344 | 2022-09-22T10:52:24.49+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04242 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:55:42.747+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04242 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:58:23.257+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04242 | 88 | AFE Power Consumption data | 2977.533 | 2022-09-22T10:50:32.917+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03579 | 41 | Barometer Temperature | -999 | 2022-09-22T10:09:50.607+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03579 | 41 | Barometer Temperature | -999 | 2022-09-22T10:41:42.72+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03579 | 41 | Barometer Temperature | -999 | 2022-09-22T11:05:14.653+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03579 | 41 | Barometer Temperature | -999 | 2022-09-23T00:19:13.427+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05105 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:17:22.823+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05105 | 88 | AFE Power Consumption data | 2587.728 | 2022-09-22T10:32:24.86+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05029 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T10:10:49.027+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05029 | 46 | Temperature Sensor Temperature | 24.945 | 2022-09-22T10:43:35.567+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03744 | 41 | Barometer Temperature | -999 | 2022-09-22T10:02:46.183+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03744 | 41 | Barometer Temperature | 24.52344 | 2022-09-22T10:40:07.723+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04866 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T09:44:16.533+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04866 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T09:47:06.913+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04866 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T09:55:25.547+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04866 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T11:17:30.777+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04866 | 46 | Temperature Sensor Temperature | -999 | 2022-09-23T00:34:53.837+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04866 | 46 | Temperature Sensor Temperature | -999 | 2022-09-23T00:36:49.577+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03938 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:03:08.823+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03938 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:05:03.713+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03938 | 88 | AFE Power Consumption data | 2501.983 | 2022-09-23T00:33:05.08+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04653 | 41 | Barometer Temperature | -999 | 2022-09-22T07:18:25.817+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04653 | 41 | Barometer Temperature | 24.32813 | 2022-09-22T07:26:20.273+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04957 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T03:39:50.897+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04957 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T05:18:19.363+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04957 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T06:25:59.57+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04957 | 46 | Temperature Sensor Temperature | -999 | 2022-09-23T00:53:55.367+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04957 | 46 | Temperature Sensor Temperature | 23.82 | 2022-09-23T03:33:44.153+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04101 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:15:51.85+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04101 | 88 | AFE Power Consumption data | 3091.857 | 2022-09-22T06:18:57.683+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04771 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:40:21.79+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04771 | 88 | AFE Power Consumption data | 2528.424 | 2022-09-22T06:08:36.957+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04593 | 41 | Barometer Temperature | -999 | 2022-09-22T01:16:29.9+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04593 | 41 | Barometer Temperature | 24.64063 | 2022-09-22T04:25:58.68+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03852 | 41 | Barometer Temperature | -999 | 2022-09-22T02:37:04.237+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03852 | 41 | Barometer Temperature | -999 | 2022-09-22T02:40:52.083+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03852 | 41 | Barometer Temperature | -999 | 2022-09-22T06:37:55.967+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03852 | 41 | Barometer Temperature | -999 | 2022-09-23T00:38:59.053+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03852 | 41 | Barometer Temperature | 23.63281 | 2022-09-23T03:32:59.793+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04960 | 88 | AFE Power Consumption data | -999 | 2022-09-22T02:16:05.24+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04960 | 88 | AFE Power Consumption data | -999 | 2022-09-22T02:20:08.523+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04960 | 41 | Barometer Temperature | -999 | 2022-09-22T06:43:06.647+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04960 | 41 | Barometer Temperature | -999 | 2022-09-23T00:39:59.07+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04321 | 41 | Barometer Temperature | -999 | 2022-09-22T01:50:30.853+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04321 | 41 | Barometer Temperature | 24.78125 | 2022-09-22T01:54:32.97+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01013 | 3 | Fixture ID | -999 | 2022-09-21T09:54:50.267+00:00 | T1_5PR1AT05
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01013 | 4 | USB Connection 1 |  | 2022-09-21T09:59:20.397+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01013 | 4 | USB Connection 1 | 1 | 2022-09-21T10:03:52.3+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04014 | 41 | Barometer Temperature | -999 | 2022-09-22T10:02:35.857+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04014 | 41 | Barometer Temperature | -999 | 2022-09-22T10:38:05.107+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04014 | 41 | Barometer Temperature | -999 | 2022-09-22T11:05:32.157+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04014 | 41 | Barometer Temperature | 22.09375 | 2022-09-23T00:14:17.867+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04873 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:59:25.77+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04873 | 88 | AFE Power Consumption data | -999 | 2022-09-22T09:01:19.38+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04873 | 88 | AFE Power Consumption data | 2549.309 | 2022-09-22T10:58:43.213+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03980 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:13:51.197+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03980 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:16:20.35+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03980 | 88 | AFE Power Consumption data | 3337.981 | 2022-09-22T10:49:25.287+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05296 | 41 | Barometer Temperature | -999 | 2022-09-22T10:19:47.583+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05296 | 41 | Barometer Temperature | 24.40625 | 2022-09-22T10:37:06.597+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04429 | 77 | D110 – 630nm Brightness (20mA) Measure diff with V_ref | 2.43 | 2022-09-22T10:24:15.75+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04429 | 77 | D110 – 630nm Brightness (20mA) Measure diff with V_ref | 804.063 | 2022-09-22T10:34:49.037+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04913 | 41 | Barometer Temperature | 47.30469 | 2022-09-22T10:14:24.077+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04913 | 41 | Barometer Temperature | 24.89844 | 2022-09-22T10:33:00.987+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03982 | 41 | Barometer Temperature | -999 | 2022-09-22T10:26:40.547+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03982 | 41 | Barometer Temperature | 24.77344 | 2022-09-22T10:30:19.077+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03762 | 41 | Barometer Temperature | -999 | 2022-09-22T08:50:27.173+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03762 | 41 | Barometer Temperature | -999 | 2022-09-22T08:56:09.993+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03762 | 41 | Barometer Temperature | 24.90625 | 2022-09-22T10:57:06.793+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04471 | 41 | Barometer Temperature | -999 | 2022-09-22T08:37:18.39+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04471 | 41 | Barometer Temperature | -999 | 2022-09-22T08:41:11.263+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04471 | 41 | Barometer Temperature | 24.24219 | 2022-09-22T11:15:56.717+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04987 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:30:40.48+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04987 | 88 | AFE Power Consumption data | 3279.216 | 2022-09-22T08:33:33.447+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03971 | 41 | Barometer Temperature | -999 | 2022-09-22T08:26:14.973+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03971 | 41 | Barometer Temperature | -999 | 2022-09-22T08:27:42.817+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03971 | 41 | Barometer Temperature | 22.88281 | 2022-09-23T00:39:15.66+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05171 | 41 | Barometer Temperature | -999 | 2022-09-22T07:33:40.337+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05171 | 41 | Barometer Temperature | -999 | 2022-09-22T07:36:32.657+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05171 | 41 | Barometer Temperature | 24.57813 | 2022-09-22T11:17:33.207+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04469 | 41 | Barometer Temperature | -999 | 2022-09-22T07:16:08.937+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04469 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:22:30.183+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04469 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:44:33.093+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04469 | 88 | AFE Power Consumption data | 3076.663 | 2022-09-23T01:02:07.007+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04946 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:56:19.937+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04946 | 41 | Barometer Temperature | -999 | 2022-09-22T04:28:12.193+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04946 | 41 | Barometer Temperature | 24.84375 | 2022-09-22T06:52:01.413+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04000 | 41 | Barometer Temperature | -999 | 2022-09-22T01:29:04.88+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04000 | 41 | Barometer Temperature | -999 | 2022-09-22T04:06:35.537+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04000 | 41 | Barometer Temperature | 24.85156 | 2022-09-22T06:49:08.217+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04310 | 88 | AFE Power Consumption data | -999 | 2022-09-22T02:55:42.027+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04310 | 68 | D101 PD dark noise OHR dev | 224.6734 | 2022-09-22T03:05:04.07+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04310 | 70 | D103 PD dark noise OHR dev | 202.439 | 2022-09-22T03:59:21.433+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04310 | 70 | D103 PD dark noise OHR dev | 196.5918 | 2022-09-22T06:39:16.13+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03693 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:49:28.28+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03693 | 88 | AFE Power Consumption data | 2443.152 | 2022-09-22T04:08:46.707+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04028 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:16:41.42+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04028 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T05:31:26.863+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04028 | 46 | Temperature Sensor Temperature | 25.235 | 2022-09-22T06:27:03.707+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04547 | 25 | D109 (660nm) Brightness and LED Current (20mA) | 0.012 | 2022-09-22T01:27:46.193+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04547 | 41 | Barometer Temperature | -999 | 2022-09-22T05:01:44.827+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04547 | 41 | Barometer Temperature | 24.6875 | 2022-09-22T06:54:02.917+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04035 | 4 | USB Connection 1 | 1 | 2022-09-21T07:07:35.87+00:00 | T1_5PR1AT06 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04035 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:33:58.64+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04035 | 88 | AFE Power Consumption data | 3127.179 | 2022-09-22T06:41:57.15+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04702 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:07:50.653+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04702 | 88 | AFE Power Consumption data | 3289.976 | 2022-09-22T10:43:26.827+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05120 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:04:10.967+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05120 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:42:03.4+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05120 | 88 | AFE Power Consumption data | 3049.533 | 2022-09-22T11:07:42.027+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04290 | 88 | AFE Power Consumption data | -999 | 2022-09-22T09:59:33.07+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04290 | 88 | AFE Power Consumption data | 3113.052 | 2022-09-22T10:37:29.543+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04677 | 112 | (LED rotation check) Internal Light Pollution (ILP) Ch0 | -999 | 2022-09-22T07:37:30.343+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04677 | 41 | Barometer Temperature | -999 | 2022-09-22T07:41:02.073+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04677 | 41 | Barometer Temperature | 25.21094 | 2022-09-22T11:19:21.297+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04540 | 75 | D109 – 940nm Brightness (20mA) Measure diff with V_ref | 2.795 | 2022-09-22T06:32:20.377+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04540 | 69 | D102 PD dark noise OHR dev | 223.869 | 2022-09-22T06:37:56.317+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04540 | 69 | D102 PD dark noise OHR dev | 190.8304 | 2022-09-22T10:06:35.733+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03951 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:26:40.633+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03951 | 88 | AFE Power Consumption data | 3085.312 | 2022-09-22T05:14:25.487+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05078 | 59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | -2.665 | 2022-09-22T03:49:01.19+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05078 | 59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | 891.098 | 2022-09-22T04:25:44.677+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03777 | 41 | Barometer Temperature | -999 | 2022-09-22T01:27:38.237+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03777 | 41 | Barometer Temperature | 24.85938 | 2022-09-22T04:16:52.347+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04459 | 41 | Barometer Temperature | -999 | 2022-09-22T01:21:55.83+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04459 | 41 | Barometer Temperature | 25.14063 | 2022-09-22T04:14:19.007+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04021 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:29:12.547+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04021 | 88 | AFE Power Consumption data | 3127.924 | 2022-09-22T04:08:48.65+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04788 | 76 | D110 – 540nm Brightness (20mA) Measure diff with V_ref | 2.324 | 2022-09-22T01:59:25.73+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04788 | 76 | D110 – 540nm Brightness (20mA) Measure diff with V_ref | 783.68 | 2022-09-22T02:02:08.443+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05048 | 3 | Fixture ID | -999 | 2022-09-22T09:44:17.847+00:00 | T1_5PR1AT05
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05048 | 3 | Fixture ID | 3 | 2022-09-22T09:47:34.31+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04564 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:57:32.783+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04564 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:00:27.897+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04564 | 88 | AFE Power Consumption data | -999 | 2022-09-23T00:30:59.59+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04564 | 88 | AFE Power Consumption data | 3154.546 | 2022-09-23T03:36:39.927+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04736 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:24:37.787+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04736 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:28:10.827+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04736 | 88 | AFE Power Consumption data | 3003.679 | 2022-09-22T11:16:25.797+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05040 | 73 | D109 – 630nm Brightness (20mA) Measure diff with V_ref | 2.793 | 2022-09-22T06:58:50.433+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05040 | 73 | D109 – 630nm Brightness (20mA) Measure diff with V_ref | 809.181 | 2022-09-22T07:14:35.053+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04730 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:01:50.62+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04730 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:03:56.95+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04730 | 88 | AFE Power Consumption data | 2459.659 | 2022-09-22T06:23:32.767+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04897 | 55 | D106 – 540nm Brightness (20mA) Measure diff with V_ref | 11.094 | 2022-09-22T05:35:36.88+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04897 | 55 | D106 – 540nm Brightness (20mA) Measure diff with V_ref | 929.301 | 2022-09-22T06:08:18.533+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05086 | 41 | Barometer Temperature | -999 | 2022-09-22T05:59:05.433+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05086 | 41 | Barometer Temperature | -999 | 2022-09-22T06:04:02.533+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05086 | 41 | Barometer Temperature | 24.5 | 2022-09-22T06:30:13.333+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04034 | 41 | Barometer Temperature | -999 | 2022-09-22T05:57:48.7+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04034 | 41 | Barometer Temperature | -999 | 2022-09-22T06:02:28.93+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04034 | 41 | Barometer Temperature | 24.53125 | 2022-09-22T06:24:15.257+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04531 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:56:42.117+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04531 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:59:26.177+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04531 | 88 | AFE Power Consumption data | 3193.962 | 2022-09-22T06:21:43.177+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04850 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:51:36.367+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04850 | 88 | AFE Power Consumption data | 3001.123 | 2022-09-22T05:26:14.397+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05336 | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | -2.058 | 2022-09-22T03:28:46.073+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05336 | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | 791.452 | 2022-09-22T05:17:16.98+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05127 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:43:06.357+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05127 | 88 | AFE Power Consumption data | 3212.777 | 2022-09-22T05:15:26.88+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04535 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T03:20:16.25+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04535 | 46 | Temperature Sensor Temperature | 24.16 | 2022-09-22T05:14:50.977+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03705 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:19:08.423+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03705 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:28:31.003+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03705 | 88 | AFE Power Consumption data | -999 | 2022-09-23T00:54:01.667+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03705 | 41 | Barometer Temperature | -999 | 2022-09-22T06:25:11.873+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04982 | 41 | Barometer Temperature | -999 | 2022-09-22T02:40:02.017+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04982 | 34 | AFE Power Consumption Test | 0.176 | 2022-09-22T02:46:08.593+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04982 | 34 | AFE Power Consumption Test | 0.144 | 2022-09-22T06:41:16.123+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04936 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:37:09.41+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04936 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:49:51.55+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04936 | 88 | AFE Power Consumption data | 3212.55 | 2022-09-22T01:39:08.163+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04521 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:32:08.567+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04521 | 79 | D110 – 940nm Brightness (20mA) Measure diff with V_ref | 2.449 | 2022-09-22T04:27:28.667+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04521 | 79 | D110 – 940nm Brightness (20mA) Measure diff with V_ref | 947.454 | 2022-09-22T06:51:39.397+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01423 | 88 | AFE Power Consumption data | -999 | 2022-09-21T10:53:34.41+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01423 | 88 | AFE Power Consumption data | 2984.323 | 2022-09-21T10:58:14.937+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01216 | 46 | Temperature Sensor Temperature | -999 | 2022-09-21T10:43:10.82+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01216 | 88 | AFE Power Consumption data | -999 | 2022-09-21T10:50:48.043+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01216 | 88 | AFE Power Consumption data | 3169.646 | 2022-09-22T11:01:02.35+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03649 | 41 | Barometer Temperature | -999 | 2022-09-22T10:14:49.32+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03649 | 41 | Barometer Temperature | -999 | 2022-09-22T10:30:09.423+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03649 | 41 | Barometer Temperature | 24.36719 | 2022-09-22T11:04:14.91+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00861 | 88 | AFE Power Consumption data | -999 | 2022-09-21T10:41:32.867+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00861 | 59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | 1.062 | 2022-09-21T10:44:40.323+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00861 | 59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | 878.892 | 2022-09-22T11:01:04.48+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04107 | 41 | Barometer Temperature | -999 | 2022-09-22T08:29:51.107+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04107 | 41 | Barometer Temperature | 24.57031 | 2022-09-22T08:33:25.43+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03746 | 41 | Barometer Temperature | -999 | 2022-09-22T07:19:30.867+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03746 | 41 | Barometer Temperature | -999 | 2022-09-22T07:23:45.747+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03746 | 41 | Barometer Temperature | 24.71875 | 2022-09-22T10:46:18.577+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04012 | 68 | D101 PD dark noise OHR dev | 208.4635 | 2022-09-22T01:33:16.347+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04012 | 68 | D101 PD dark noise OHR dev | 229.4114 | 2022-09-22T06:48:56.377+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04012 | 68 | D101 PD dark noise OHR dev | 220.4895 | 2022-09-22T10:53:40.793+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04012 | 68 | D101 PD dark noise OHR dev | 182.6931 | 2022-09-23T00:50:49.787+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04012 | 70 | D103 PD dark noise OHR dev | 203.6565 | 2022-09-22T04:08:43.443+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04770 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:47:22.947+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04770 | 88 | AFE Power Consumption data | 3217.426 | 2022-09-22T05:50:50.927+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04144 | 41 | Barometer Temperature | -999 | 2022-09-22T01:15:44.81+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04144 | 41 | Barometer Temperature | -999 | 2022-09-22T04:23:47.287+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04144 | 41 | Barometer Temperature | 24.48438 | 2022-09-22T06:52:12.407+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04327 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:16:31.193+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04327 | 34 | AFE Power Consumption Test | 0.177 | 2022-09-22T04:03:58.613+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04327 | 34 | AFE Power Consumption Test | 0.15 | 2022-09-22T06:46:19.95+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03895 | 59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | -2.383 | 2022-09-22T02:42:10.52+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03895 | 59 | D108 – 540nm Brightness (20mA) Measure diff with V_ref | 935.803 | 2022-09-22T02:49:30.623+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04583 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:54:25.19+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04583 | 88 | AFE Power Consumption data | 3098.427 | 2022-09-22T01:57:24.34+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04942 | 41 | Barometer Temperature | -999 | 2022-09-22T00:34:31.327+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04942 | 41 | Barometer Temperature | -999 | 2022-09-23T01:02:17.84+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04942 | 41 | Barometer Temperature | 23.80469 | 2022-09-23T03:33:32.71+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04942 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:21:40.237+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04942 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:17:53.333+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04942 | 88 | AFE Power Consumption data | -999 | 2022-09-23T00:56:45.34+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04258 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:46:10.33+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04258 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:48:31.307+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04258 | 88 | AFE Power Consumption data | 3076.191 | 2022-09-22T11:02:10.087+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05277 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:13:46.037+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05277 | 88 | AFE Power Consumption data | 3034.591 | 2022-09-22T10:39:01.473+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04119 | 41 | Barometer Temperature | -999 | 2022-09-22T10:25:32.56+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04119 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:28:39.89+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04119 | 88 | AFE Power Consumption data | 3026.206 | 2022-09-22T11:08:22.38+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04528 | 78 | D110 – 660nm Brightness (20mA) Measure diff with V_ref | 2.458 | 2022-09-22T07:32:34.483+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04528 | 78 | D110 – 660nm Brightness (20mA) Measure diff with V_ref | 970.099 | 2022-09-22T07:35:22.1+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04513 | 41 | Barometer Temperature | -999 | 2022-09-22T06:09:56.99+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04513 | 41 | Barometer Temperature | 24.09375 | 2022-09-22T06:14:16.977+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04507 | 41 | Barometer Temperature | -999 | 2022-09-22T01:23:17.73+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04507 | 41 | Barometer Temperature | 24.95313 | 2022-09-22T04:29:23.783+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04520 | 41 | Barometer Temperature | -999 | 2022-09-22T01:19:33.24+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04520 | 41 | Barometer Temperature | 24.35156 | 2022-09-22T04:25:17.053+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05124 | 41 | Barometer Temperature | -999 | 2022-09-22T03:31:22.767+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05124 | 41 | Barometer Temperature | 23.85938 | 2022-09-22T05:18:40.49+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04577 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:52:10.833+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04577 | 88 | AFE Power Consumption data | 3095.675 | 2022-09-22T01:54:52.033+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00953 | 104 | D110 – 630nm Brightness (20mA) | -999 | 2022-09-22T01:44:30.08+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00953 | 104 | D110 – 630nm Brightness (20mA) | 1 | 2022-09-22T01:49:37.423+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04457 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:23:03.393+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04457 | 88 | AFE Power Consumption data | 2938.616 | 2022-09-22T04:17:16.337+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01005 | 49 | AFE Part ID Read | -999 | 2022-09-21T10:14:50.31+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01005 | 49 | AFE Part ID Read | 808800 | 2022-09-21T10:18:58.143+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00972 | 49 | AFE Part ID Read | -999 | 2022-09-21T10:09:21.037+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00972 | 49 | AFE Part ID Read | 808800 | 2022-09-21T10:12:54.59+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04632 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:15:28.767+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04632 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:33:36.843+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04632 | 88 | AFE Power Consumption data | -999 | 2022-09-22T11:09:29.263+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04632 | 88 | AFE Power Consumption data | 2535.192 | 2022-09-23T00:27:21.523+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04872 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:00:44.147+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04872 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:14:06.077+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04872 | 88 | AFE Power Consumption data | 2477.837 | 2022-09-22T10:49:13.527+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04095 | 41 | Barometer Temperature | -999 | 2022-09-22T10:18:53.617+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04095 | 41 | Barometer Temperature | -999 | 2022-09-22T10:33:46.34+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04095 | 41 | Barometer Temperature | 24.35156 | 2022-09-22T11:13:20.793+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04662 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:18:53.227+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04662 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:32:10.563+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04662 | 88 | AFE Power Consumption data | 3295.092 | 2022-09-22T11:10:08.153+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03132 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:44:02.677+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03132 | 88 | AFE Power Consumption data | 2532.624 | 2022-09-22T08:47:17.56+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00996 | 41 | Barometer Temperature | -999 | 2022-09-21T10:18:51.62+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E00996 | 41 | Barometer Temperature | 23.91406 | 2022-09-21T10:22:26.393+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04773 | 41 | Barometer Temperature | -999 | 2022-09-22T10:04:01.727+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04773 | 41 | Barometer Temperature | -999 | 2022-09-22T10:40:01.67+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04773 | 41 | Barometer Temperature | 24.49219 | 2022-09-22T11:10:23.633+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03271 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:15:29.967+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03271 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:36:05.863+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03271 | 79 | D110 – 940nm Brightness (20mA) Measure diff with V_ref | 2.54 | 2022-09-22T11:07:55.41+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03271 | 79 | D110 – 940nm Brightness (20mA) Measure diff with V_ref | 958.512 | 2022-09-23T00:33:47.587+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05259 | 88 | AFE Power Consumption data | -999 | 2022-09-22T07:39:08.863+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05259 | 88 | AFE Power Consumption data | 3034.321 | 2022-09-22T07:43:54.867+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E02667 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:42:10.897+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E02667 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:28:39.593+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E02667 | 88 | AFE Power Consumption data | 3028.286 | 2022-09-22T06:28:33.24+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03981 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:42:55.58+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03981 | 88 | AFE Power Consumption data | 3183.259 | 2022-09-22T05:47:44.947+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05074 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T03:42:02.867+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05074 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:24:46.51+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05074 | 88 | AFE Power Consumption data | -999 | 2022-09-22T06:20:16.057+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05074 | 88 | AFE Power Consumption data | 2489.225 | 2022-09-23T00:57:16.867+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05051 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:58:27.05+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05051 | 88 | AFE Power Consumption data | 2941.448 | 2022-09-22T04:03:35.787+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04688 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:22:18.72+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04688 | 88 | AFE Power Consumption data | 2546.991 | 2022-09-22T05:17:20.583+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04243 | 71 | D104 PD dark noise OHR dev | 203.4859 | 2022-09-22T00:45:55.033+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04243 | 71 | D104 PD dark noise OHR dev | 182.2883 | 2022-09-22T01:38:09.687+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01267 | 21 | D109 (540nm) Brightness and LED Current (20mA) | 0.012 | 2022-09-22T01:28:26.527+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E01267 | 21 | D109 (540nm) Brightness and LED Current (20mA) | 20.57 | 2022-09-22T01:32:32.853+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04661 | 62 | D103 PD dark noise SpO2 dev | 98.0736 | 2022-09-22T00:12:41.1+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04661 | 71 | D104 PD dark noise OHR dev | 223.2083 | 2022-09-22T05:30:15.467+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04661 | 60 | D101 PD dark noise SpO2 dev | 97.66671 | 2022-09-22T06:24:16.067+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04661 | 41 | Barometer Temperature | -999 | 2022-09-23T00:45:03.967+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04661 | 49 | AFE Part ID Read | -999 | 2022-09-23T00:47:18.603+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04959 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:14:15.687+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04959 | 49 | AFE Part ID Read | -999 | 2022-09-22T10:40:12.613+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04959 | 49 | AFE Part ID Read | 808800 | 2022-09-22T11:13:39.523+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05012 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:43:59.657+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05012 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:48:57.397+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05012 | 88 | AFE Power Consumption data | 3054.209 | 2022-09-22T11:15:06.453+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04169 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:35:39.603+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04169 | 88 | AFE Power Consumption data | 3076.249 | 2022-09-22T06:07:56.787+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04384 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:53:11.48+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04384 | 88 | AFE Power Consumption data | -999 | 2022-09-22T05:57:33.977+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04384 | 88 | AFE Power Consumption data | 2935.014 | 2022-09-22T06:21:23.973+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04916 | 10 | D103 Response avg | -999 | 2022-09-22T03:29:44.457+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04916 | 10 | D103 Response avg | 20531.44 | 2022-09-22T05:05:37.58+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03946 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:27:47.08+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03946 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:31:26.963+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03946 | 88 | AFE Power Consumption data | 3005.959 | 2022-09-23T00:37:53.497+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05228 | 76 | D110 – 540nm Brightness (20mA) Measure diff with V_ref | 2.414 | 2022-09-22T07:28:06+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05228 | 76 | D110 – 540nm Brightness (20mA) Measure diff with V_ref | 786.916 | 2022-09-22T07:32:44.71+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04354 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:14:21.807+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04354 | 88 | AFE Power Consumption data | -999 | 2022-09-22T04:22:23.693+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04354 | 88 | AFE Power Consumption data | 2403.277 | 2022-09-22T06:49:28.107+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05129 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:27:35.917+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05129 | 88 | AFE Power Consumption data | 2527.542 | 2022-09-22T05:25:49.01+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03976 | 46 | Temperature Sensor Temperature | -999 | 2022-09-22T01:21:24.68+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03976 | 46 | Temperature Sensor Temperature | 25.33 | 2022-09-22T04:29:29.61+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05119 | 41 | Barometer Temperature | -999 | 2022-09-22T02:28:36.257+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05119 | 41 | Barometer Temperature | 24.53125 | 2022-09-22T02:33:40.78+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04664 | 88 | AFE Power Consumption data | -999 | 2022-09-22T10:20:47.37+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04664 | 88 | AFE Power Consumption data | 2884.491 | 2022-09-22T10:32:36.38+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04217 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:52:14.2+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04217 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:58:36.677+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04217 | 88 | AFE Power Consumption data | 3213.833 | 2022-09-22T10:58:16.983+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04646 | 41 | Barometer Temperature | -999 | 2022-09-22T00:54:24.807+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04646 | 41 | Barometer Temperature | 25.14063 | 2022-09-22T04:27:12.033+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03876 | 88 | AFE Power Consumption data | -999 | 2022-09-22T09:34:37.897+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03876 | 88 | AFE Power Consumption data | 2977.519 | 2022-09-22T09:37:27.083+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05285 | 41 | Barometer Temperature | -999 | 2022-09-22T03:28:39.997+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E05285 | 41 | Barometer Temperature | 24.17969 | 2022-09-22T05:15:09.547+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04489 | 88 | AFE Power Consumption data | -999 | 2022-09-22T08:08:59.2+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04489 | 88 | AFE Power Consumption data | 2382.423 | 2022-09-22T08:15:03.89+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04953 | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | 2.893 | 2022-09-22T02:36:03.427+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04953 | 81 | D109 – 540nm Brightness (20mA) Measure diff with V_ref | 791.324 | 2022-09-22T02:44:42.7+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03883 | 88 | AFE Power Consumption data | -999 | 2022-09-22T01:26:44.98+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E03883 | 88 | AFE Power Consumption data | 2500.238 | 2022-09-22T04:30:27.327+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04839 | 88 | AFE Power Consumption data | -999 | 2022-09-22T00:52:52.427+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04839 | 88 | AFE Power Consumption data | 2473.145 | 2022-09-22T05:23:01.927+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04992 | 88 | AFE Power Consumption data | -999 | 2022-09-22T03:14:23.71+00:00 | T1_5PR1AT05 | 3
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04992 | 88 | AFE Power Consumption data | 2414.864 | 2022-09-22T05:29:01.883+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04134 | 77 | D110 – 630nm Brightness (20mA) Measure diff with V_ref | 1.572 | 2022-09-22T02:28:31.993+00:00 | T1_5PR1AT05 | 2
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04134 | 77 | D110 – 630nm Brightness (20mA) Measure diff with V_ref | 783.931 | 2022-09-22T02:36:13.887+00:00 | T1_5PR1AT05 | 1
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04742 | 41 | Barometer Temperature | -999 | 2022-09-22T02:27:11.273+00:00 | T1_5PR1AT05 | 4
16045256 | 17859 | Epix_Pro_ATE_TI | .152 | 2905E04742 | 41 | Barometer Temperature | 24.60156 | 2022-09-22T02:30:28.04+00:00 | T1_5PR1AT05 | 1
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04723 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T00:12:05.347+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04723 | 30 | Motor Resonant Frequency | 185 | 2022-09-22T00:17:34.517+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04557 | 30 | Motor Resonant Frequency | -999 | 2022-09-21T08:27:58.627+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04557 | 30 | Motor Resonant Frequency | 186 | 2022-09-21T08:30:14.803+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04597 | 20 | ECG 2mv 5Hz Test | 1.873398 | 2022-09-21T08:27:34.637+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04597 | 20 | ECG 2mv 5Hz Test | 1.880646 | 2022-09-21T08:38:54.79+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04597 | 20 | ECG 2mv 5Hz Test | 1.846313 | 2022-09-22T09:51:19.7+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04597 | 59 | ECG 2mv 2Hz Test | 1.817322 | 2022-09-21T08:33:17.46+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04814 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T01:14:14.61+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04814 | 30 | Motor Resonant Frequency | 187 | 2022-09-22T01:17:18.823+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04732 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T01:12:25.467+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04732 | 30 | Motor Resonant Frequency | 187 | 2022-09-22T01:15:02.833+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04354 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T10:11:55.677+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04354 | 30 | Motor Resonant Frequency | 187 | 2022-09-21T10:14:10.143+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E03774 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T07:21:06.697+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E03774 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T07:23:26.643+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E03774 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T08:20:03.13+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04730 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T05:42:12.017+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04730 | 30 | Motor Resonant Frequency | 188 | 2022-09-22T05:51:44.317+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05007 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:27:42.947+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05007 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:30:38.143+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05007 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:33:59.11+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04957 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T01:28:18.053+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04957 | 30 | Motor Resonant Frequency | 188 | 2022-09-22T01:30:41.78+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05019 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T07:49:46.29+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05019 | 30 | Motor Resonant Frequency | 191 | 2022-09-22T08:17:40.703+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05225 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:26:52.357+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05225 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:30:01.74+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05225 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:34:27.68+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04960 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T00:13:36.753+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04960 | 30 | Motor Resonant Frequency | 191 | 2022-09-22T00:18:09.467+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04310 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T08:32:56.097+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04310 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T08:38:27.067+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04524 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T08:27:42.197+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04524 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T08:29:59.913+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 20 | ECG 2mv 5Hz Test | 0.309753 | 2022-09-22T02:39:30.31+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 20 | ECG 2mv 5Hz Test | 0.390244 | 2022-09-22T02:42:28.057+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 20 | ECG 2mv 5Hz Test | 0.315475 | 2022-09-22T02:57:15.447+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 20 | ECG 2mv 5Hz Test | 0.312042 | 2022-09-22T09:48:25.323+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 9 | Is ESN check in |  | 2022-09-22T02:44:54.233+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 9 | Is ESN check in |  | 2022-09-22T02:47:33.507+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 9 | Is ESN check in |  | 2022-09-22T02:50:45.16+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04159 | 27 | ECG 2mv 1Hz Test | 0.318146 | 2022-09-22T03:00:56.78+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04045 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T02:13:06.157+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04045 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T02:25:43.503+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04045 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T02:29:37.933+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04321 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T09:38:38.163+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04321 | 30 | Motor Resonant Frequency | 193 | 2022-09-21T09:41:36.687+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04434 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T07:41:04.343+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04434 | 30 | Motor Resonant Frequency | 190 | 2022-09-22T08:12:38.23+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05053 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:17:33.657+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05053 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:20:33.643+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05053 | 30 | Motor Resonant Frequency | -999 | 2022-09-22T01:27:15.367+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04848 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T00:38:12.59+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04848 | 30 | Motor Resonant Frequency | 187 | 2022-09-22T00:42:12.937+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04349 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T10:05:01.823+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04349 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T10:07:16.427+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E03777 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T09:55:35.987+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E03777 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T09:57:56.823+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04241 | 30 | Motor Resonant Frequency | 182 | 2022-09-21T09:43:11.953+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04241 | 30 | Motor Resonant Frequency | 188 | 2022-09-21T09:45:45.187+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04119 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T02:19:24.313+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04119 | 30 | Motor Resonant Frequency | 189 | 2022-09-22T10:23:51.407+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04093 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T08:43:21.787+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04093 | 30 | Motor Resonant Frequency | 187 | 2022-09-21T08:46:07.407+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04679 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T00:40:27.893+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04679 | 30 | Motor Resonant Frequency | 194 | 2022-09-22T00:41:31.913+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04679 | 30 | Motor Resonant Frequency | 187 | 2022-09-22T00:44:33.367+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04647 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T10:42:46.273+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04647 | 30 | Motor Resonant Frequency | 189 | 2022-09-21T10:45:11.25+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04904 | 15 | Check 1907IOP | -999 | 2022-09-22T00:59:00.713+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04904 | 15 | Check 1907IOP | -999 | 2022-09-22T01:02:15.7+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04904 | 15 | Check 1907IOP | -999 | 2022-09-22T01:09:06.98+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04600 | 15 | Check 1907IOP | -999 | 2022-09-21T10:26:00.293+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04600 | 15 | Check 1907IOP | -999 | 2022-09-21T10:33:42.27+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E00861 | 59 | ECG 2mv 2Hz Test | -999 | 2022-09-21T08:01:07.36+00:00 | T1_5PR1AT06 | 3
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E00861 | 59 | ECG 2mv 2Hz Test | 2.191925 | 2022-09-21T08:03:58.017+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E03888 | 7 | USB Connection (1) |  | 2022-09-22T07:29:19.123+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E03888 | 7 | USB Connection (1) | 1 | 2022-09-22T07:31:37.24+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04054 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T00:18:12.337+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04054 | 30 | Motor Resonant Frequency | 188 | 2022-09-22T00:20:34.7+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04566 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T10:14:25.18+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04566 | 30 | Motor Resonant Frequency | 189 | 2022-09-21T10:16:56.777+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04471 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T07:38:35.317+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04471 | 30 | Motor Resonant Frequency | 191 | 2022-09-22T08:15:32.803+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04139 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T01:21:11.683+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04139 | 30 | Motor Resonant Frequency | 192 | 2022-09-22T01:25:18.387+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04593 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T10:09:49.163+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04593 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T10:12:07.89+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04437 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T08:54:37.1+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04437 | 30 | Motor Resonant Frequency | 188 | 2022-09-21T09:33:07.753+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E00624 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T08:05:36.613+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E00624 | 30 | Motor Resonant Frequency | 189 | 2022-09-21T08:07:54.587+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04030 | 30 | Motor Resonant Frequency | -999 | 2022-09-21T09:36:17.023+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04030 | 30 | Motor Resonant Frequency | -999 | 2022-09-21T09:38:18.46+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04030 | 30 | Motor Resonant Frequency | -999 | 2022-09-21T09:43:24.457+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04786 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T10:24:40.12+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04786 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T10:31:26.557+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04487 | 30 | Motor Resonant Frequency | 183 | 2022-09-21T10:16:53.997+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04487 | 30 | Motor Resonant Frequency | 189 | 2022-09-21T10:19:38.143+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04286 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T02:47:41.877+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04286 | 30 | Motor Resonant Frequency | 189 | 2022-09-22T02:56:55.563+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04446 | 30 | Motor Resonant Frequency | -999 | 2022-09-21T10:17:05.607+00:00 | T1_5PR1AT06 | 4
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04446 | 30 | Motor Resonant Frequency | -999 | 2022-09-21T10:21:27.233+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E01100 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T08:15:00.353+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E01100 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T08:17:00.083+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05228 | 30 | Motor Resonant Frequency | 184 | 2022-09-22T06:51:15.033+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E05228 | 30 | Motor Resonant Frequency | 188 | 2022-09-22T06:53:35.987+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04036 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T08:45:18.153+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04036 | 30 | Motor Resonant Frequency | 191 | 2022-09-21T08:48:34.36+00:00 | T1_5PR1AT06 | 2
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04831 | 30 | Motor Resonant Frequency | 184 | 2022-09-21T08:43:32.573+00:00 | T1_5PR1AT06 | 255
16045256 | 17858 | Epix_Pro_ECG_ATE | .704 | 2905E04831 | 30 | Motor Resonant Frequency | 190 | 2022-09-21T08:50:45.783+00:00 | T1_5PR1AT06 | 2